Instruction Sheet for the Candidate

| Qualification | National Vocational Certificate-Level-3 in Water Quality & Resource Management (Water Quality Surveyor) |
| --- | --- |
| Competency Standard | Testing of Electrical Conductivity (EC) by EC Meter |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name______________________________________________________ Registration/Roll Number_______________________________________ |
| Guidance for Candidate | To meet this standard, you are required to complete the following within 04 Hrs. time frame (for practical demonstration & assessment): - Prepare samples for testing - Prepare for laboratory testing - Perform tests on samples - Apply Quality Control Checks - Record the results/ Finalize work - Adopt precautions during work |
| Time: 04 Hrs. | During a practical assessment, under observation by an assessor, you are required to |

| Minimum Evidence Required | Prepare samples for testing - Check the sample label for the requirement of EC testing and maintain the Laboratory room temperature at 25 °C through heating or cooling whatever required. - Check for the availability of EC standard of concentration 1413 µS/cm. - Identify appropriate test method, equipment and safety requirements. Prepare for laboratory testing - Set up EC meter and/or reagents in accordance with the standard work instructions. - Conduct pre-use and safety checks. - Check the calibration status of equipment. Perform calibration if required. Perform tests on samples - Turn on the instrument and give it 5 minutes warm up time. Press the desired mode of measurement. - check the calibration by running known EC standard of concentration 1413 µS/cm. - Perform test by taking 50 ml deionized water in a beaker, immerse the probe in it and stir the blank until the instrument gives stable EC reading. - Analyze three replicates (3 readings per sample) - Store unused reagents and dispose of wastes as required by relevant regulations and codes. - Clean and store equipment Apply Quality Control Checks - standardize electrical Conductivity meter by running known standard. - Analyze method blank after every 10 samples. - Analyze Laboratory Control samples of 1413 µS/cm analyze with every batch of 10 samples. - check repeatability by reanalyzing a pre-analyzed sample once a day - Record Daily quality control data in QC Data sheet-I. Record the results/ Finalize work - Note down the results on analyst workbook. - Report the results on result record form with proper uncertainty and unit. Adopt precautions during work - Check calibration. - Adopt precautionary measurements for probe handling. |
| --- | --- |

Self-Assessment Checklist

| Candidate Name |  |
| --- | --- |
| Registration No. |  |
| Qualification | National Vocational Certificate-Level-3 in Water Quality & Resource Management (Water Quality Surveyor) |
| Competency Standard | Testing of Electrical Conductivity (EC) by EC Meter |
| Purpose of Assessment | Formative Assessment |
| Assessment Task | - Prepare samples for testing - Prepare for laboratory testing - Perform tests on samples - Apply Quality Control Checks - Record the results/ Finalize work - Adopt precautions during work |

I can……………….

| Performance Criteria | Yes | No |
| --- | --- | --- |
| - Check the sample label for the requirement of EC testing and maintain the Laboratory room temperature at 25 °C through heating or cooling whatever required. |  |  |
| - Check for the availability of EC standard of concentration 1413 µS/cm. |  |  |
| - Set up EC meter and/or reagents in accordance with the standard work instructions. |  |  |
| - Conduct pre-use and safety checks. |  |  |
| - Check the calibration status of equipment. Perform calibration if required. |  |  |
| - Check the calibration status of equipment. Perform calibration if required. |  |  |
| - Turn on the instrument and give it 5 minutes warm up time. Press the desired mode of measurement. |  |  |
| - check the calibration by running known EC standard of concentration 1413 µS/cm. |  |  |
| - Perform test by taking 50 ml deionized water in a beaker, immerse the probe in it and stir the blank until the instrument gives stable EC reading. |  |  |
| - Analyze three replicates (3 readings per sample) |  |  |
| - Store unused reagents and dispose of wastes as required by relevant regulations and codes. |  |  |
| - Clean and store equipment |  |  |
| - standardize electrical Conductivity meter by running known standard. |  |  |
| - Analyze method blank after every 10 samples. |  |  |
| - Analyze Laboratory Control samples of 1413 µS/cm analyze with every batch of 10 samples. |  |  |
| - check repeatability by reanalyzing a pre-analyzed sample once a day |  |  |
| - Note down the results on analyst workbook. |  |  |
| - Report the results on result record form with proper uncertainty and unit. |  |  |
| - Check calibration. |  |  |
| - Adopt precautionary measurements for probe handling. |  |  |

Candidate’s Signature__________________ Assessor’s Signature____________________

Date: _______________________________

Assessors Judgment Guide

| Qualification | National Vocational Certificate-Level-3 in Water Quality & Resource Management (Water Quality Surveyor) |
| --- | --- |
| Competency Standard | Testing of Electrical Conductivity (EC) by EC Meter |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: ______________________ Signature: _______________ |
| Assessment Outcome | COMPETENT NOT YET COMPETENT Name of the Assessor________________________________________________ Assessor’s code: ____________________________________________________ Signature: ________________________________ |

| Assessment Summary (to be filled by the assessor) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Activity | Method |  |  |  |  | Result |  |
| Nature of Activity | Written | Oral | Observation | Portfolio | Role Play | Competent | Not Yet Competent |
| Practical Skill Demonstration |  |  |  |  |  |  |  |
| Knowledge Assessment |  |  |  |  |  |  |  |
| Other Requirement |  |  |  |  |  |  |  |

Observation Checklist

| Assessment Task |  | - Prepare samples for testing - Prepare for laboratory testing - Perform tests on samples - Apply Quality Control Checks - Record the results/ Finalize work - Adopt precautions during work |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| During the practical assessment, candidate demonstrated the following: |  |  |  | Yes | No | Remarks |
|  | Check the sample label for the requirement of EC testing and maintain the Laboratory room temperature at 25 °C through heating or cooling whatever required. |  |  |  |  |  |
|  | Check for the availability of EC standard of concentration 1413 µS/cm. |  |  |  |  |  |
|  | Set up EC meter and/or reagents in accordance with the standard work instructions. |  |  |  |  |  |
|  | Conduct pre-use and safety checks. |  |  |  |  |  |
|  | Check the calibration status of equipment. Perform calibration if required. |  |  |  |  |  |
|  | Check the calibration status of equipment. Perform calibration if required. |  |  |  |  |  |
|  | Turn on the instrument and give it 5 minutes warm up time. Press the desired mode of measurement. |  |  |  |  |  |
|  | check the calibration by running known EC standard of concentration 1413 µS/cm. |  |  |  |  |  |
|  | Perform test by taking 50 ml deionized water in a beaker, immerse the probe in it and stir the blank until the instrument gives stable EC reading. |  |  |  |  |  |
|  | Analyze three replicates (3 readings per sample) |  |  |  |  |  |
|  | Store unused reagents and dispose of wastes as required by relevant regulations and codes. |  |  |  |  |  |
|  | Clean and store equipment |  |  |  |  |  |
|  | standardize electrical Conductivity meter by running known standard. |  |  |  |  |  |
|  | Analyze method blank after every 10 samples. |  |  |  |  |  |
|  | Analyze Laboratory Control samples of 1413 µS/cm analyze with every batch of 10 samples. |  |  |  |  |  |
|  | check repeatability by reanalyzing a pre-analyzed sample once a day |  |  |  |  |  |
|  | Record Daily quality control data in QC Data sheet-I. |  |  |  |  |  |
|  | Report the results on result record form with proper uncertainty and unit. |  |  |  |  |  |
|  | Check calibration. |  |  |  |  |  |
|  | Adopt precautionary measurements for probe handling. |  |  |  |  |  |
| Competent |  |  | Not Yet Competent |  |  |  |

Knowledge Assessment

| Qualification | National Vocational Certificate-Level-3 in Water Quality & Resource Management (Water Quality Surveyor) |
| --- | --- |
| Competency Standard | Testing of Electrical Conductivity (EC) by EC Meter |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: _________________ Candidate Signature: ___________ |
| Assessment Outcome | COMPETENT NOT YET COMPETENT Name of the Assessor: ________________________________________________ Assessor’s code: ____________________________________________________ Signature of the Assessor: _______________________ |

Candidate’s response is not required to be identical, but similar concepts and/or keywords must be used. Oral questioning may be used to clarify candidate understanding of topic and its application.

| Questions (Candidate confidently answered questions correctly and demonstrated understanding of the topics and their application) |  | Satisfactory | Not Satisfactory |
| --- | --- | --- | --- |
|  | What is the unit of electrical conductivity? |  |  |
|  | Which standard is used for calibration of EC meter? |  |  |
|  | Differentiate between mili siemens (mS/cm) and micr osiemens/ cm( µS/cm) . |  |  |
|  | What is the specification of glass electrode of EC meter? |  |  |
|  | How to prepare 0.01 M standard of Kcl ? |  |  |
|  | What is the relation between EC and TDS? |  |  |
|  | How TDS is calculated in Lab? |  |  |

Feedback to the Candidate

Candidate’s Signature__________________ Assessor’s Signature _________________