Instruction Sheet for the Candidate

| Qualification | National Vocational Certificate-Level-3 in Water Quality & Resource Management (Water Quality Surveyor) |
| --- | --- |
| Competency Standard | Testing of Turbidity by Turbidity Meter |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name______________________________________________________ Registration/Roll Number_______________________________________ |
| Guidance for Candidate | To meet this standard, you are required to complete the following within 04 Hrs. time frame (for practical demonstration & assessment): - Perform Testing of turbidity by turbidity Meter - Prepare for laboratory testing - Perform tests on samples - Quality Control Checks - Record the results/ Finalize work - Adopt precautions during result |
| Time: 04 Hrs. | During a practical assessment, under observation by an assessor, you are required to |

| Minimum Evidence Required | Perform Testing of turbidity by turbidity Meter - Check the sample label for the requirement of Turbidity testing. - Check for the availability of Turbidity standard of10 NTU, otherwise prepare from 100 NTU Stock solution through dilutions as per requirements. - Identify appropriate test method, equipment and safety requirements. Prepare for laboratory testing - Set up test equipment Turbidity meter and/or reagents in accordance with the specified work instructions. - Conduct pre-use and safety checks. - Check the calibration status of equipment. Perform calibration if required. Perform tests on samples - Turn the meter on by pressing specific key. - Rinse the cuvet with turbidity free water and shake the sample bottle gently. - Fill the cuvet up to the mark with thoroughly agitated sample. - Wipe the cuvet thoroughly with lint free tissue, to make the cuvet free of oil drift and fingerprints. - Shake the sample vigorously. - Place the cuvet into the cell, the mark on the cuvet cap should point towards the LCD/screen. - Run the sample as per work instructions - Analyze three replicates - Store unused reagents and dispose of wastes as required by relevant regulations and codes. - Clean and store equipment. Apply Quality Control Checks - Standardize Turbidity Meter by running known standards. - Analyzed Method blank after every 10 samples - Run the Laboratory Control samples of 10 FTU/NTU with every batch of 10 samples - check repeatability by reanalyzing a pre-analyzed sample once a day - Record Daily quality control data in QC Data sheet-I. Record the results/ Finalise work - Note down the Results on analyst workbook. - Calculate uncertainty of theresults - Record the results on result record form . - Clear and restore work area. Adopt precautions during work - Adopt precautionary measures for cuvette handling. - Follow Proper procedure for sample handling . - Check glassware cleanliness. Dirty glassware could also affect readings along with scratched or edged vials. |
| --- | --- |

Self-Assessment Checklist

| Candidate Name |  |
| --- | --- |
| Registration No. |  |
| Qualification | National Vocational Certificate-Level-3 in Water Quality & Resource Management (Water Quality Surveyor) |
| Competency Standard | Testing of Turbidity by Turbidity Meter |
| Purpose of Assessment | Formative Assessment |
| Assessment Task | - Perform Testing of pH by pH Meter - Prepare for laboratory testing - Perform tests on samples - Quality Control Checks - Record the results/ Finalize work - Adopt precautions during result |

I can……………….

| Performance Criteria | Yes | No |
| --- | --- | --- |
| - Check the sample label for the requirement of Turbidity testing. |  |  |
| - Check for the availability of Turbidity standard of10 NTU, otherwise prepare from 100 NTU Stock solution through dilutions as per requirements. |  |  |
| - Identify appropriate test method, equipment and safety requirements. |  |  |
| - Check the sample label for the requirement of Turbidity testing. |  |  |
| - Keep the sample at room temperature for few minutes. |  |  |
| - Check for the availability of Turbidity standard of10 NTU, otherwise prepare from 100 NTU Stock solution through dilutions as per requirements. |  |  |
| - Turn the meter on by pressing specific key. |  |  |
| - Rinse the cuvet with turbidity free water and shake the sample bottle gently. |  |  |
| - Fill the cuvet up to the mark with thoroughly agitated sample. |  |  |
| - Wipe the cuvet thoroughly with lint free tissue, to make the cuvet free of oil drift and fingerprints. |  |  |
| - Shake the sample vigorously. |  |  |
| - Place the cuvet into the cell, the mark on the cuvet cap should point towards the LCD/screen. |  |  |
| - Run the sample as per work instructions |  |  |
| - Analyze three replicates |  |  |
| - Store unused reagents and dispose of wastes as required by relevant regulations and codes. |  |  |
| - Clean and store equipment. |  |  |
| - Standardize Turbidity Meter by running known standards. |  |  |
| - Analyzed Method blank after every 10 samples |  |  |
| - Run the Laboratory Control samples of 10 FTU/NTU with every batch of 10 samples |  |  |
| - check repeatability by reanalyzing a pre-analyzed sample once a day |  |  |
| - Record Daily quality control data in QC Data sheet-I. |  |  |
| - Note down the Results on analyst workbook. |  |  |
| - Calculate uncertainty of theresults |  |  |
| - Record the results on result record form . |  |  |
| - Clear and restore work area. |  |  |
| - Adopt precautionary measures for cuvette handling. |  |  |
| - Follow Proper procedure for sample handling . |  |  |
| - Check glassware cleanliness. Dirty glassware could also affect readings along with scratched or edged vials. |  |  |

Candidate’s Signature__________________ Assessor’s Signature____________________

Date: _______________________________

Assessors Judgment Guide

| Qualification | National Vocational Certificate-Level-3 in Water Quality & Resource Management (Water Quality Surveyor) |
| --- | --- |
| Competency Standard | Testing of Turbidity by Turbidity Meter |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: ______________________ Signature: _______________ |
| Assessment Outcome | COMPETENT NOT YET COMPETENT Name of the Assessor________________________________________________ Assessor’s code: ____________________________________________________ Signature: ________________________________ |

| Assessment Summary (to be filled by the assessor) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Activity | Method |  |  |  |  | Result |  |
| Nature of Activity | Written | Oral | Observation | Portfolio | Role Play | Competent | Not Yet Competent |
| Practical Skill Demonstration |  |  |  |  |  |  |  |
| Knowledge Assessment |  |  |  |  |  |  |  |
| Other Requirement |  |  |  |  |  |  |  |

Observation Checklist

| Assessment Task |  | - Perform Testing of turbidity by turbidity Meter - Prepare for laboratory testing - Perform tests on samples - Quality Control Checks - Record the results/ Finalize work - Adopt precautions during result |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| During the practical assessment, candidate demonstrated the following: |  |  |  | Yes | No | Remarks |
|  | Check the sample label for the requirement of Turbidity testing. |  |  |  |  |  |
|  | Check for the availability of Turbidity standard of10 NTU, otherwise prepare from 100 NTU Stock solution through dilutions as per requirements. |  |  |  |  |  |
|  | Identify appropriate test method, equipment and safety requirements. |  |  |  |  |  |
|  | Check the sample label for the requirement of Turbidity testing. |  |  |  |  |  |
|  | Keep the sample at room temperature for few minutes. |  |  |  |  |  |
|  | Check for the availability of Turbidity standard of10 NTU, otherwise prepare from 100 NTU Stock solution through dilutions as per requirements. |  |  |  |  |  |
|  | Turn the meter on by pressing specific key. |  |  |  |  |  |
|  | Rinse the cuvet with turbidity free water and shake the sample bottle gently. |  |  |  |  |  |
|  | Fill the cuvet up to the mark with thoroughly agitated sample. |  |  |  |  |  |
|  | Wipe the cuvet thoroughly with lint free tissue, to make the cuvet free of oil drift and fingerprints. |  |  |  |  |  |
|  | Shake the sample vigorously. |  |  |  |  |  |
|  | Place the cuvet into the cell, the mark on the cuvet cap should point towards the LCD/screen. |  |  |  |  |  |
|  | Run the sample as per work instructions |  |  |  |  |  |
|  | Analyze three replicates |  |  |  |  |  |
|  | Store unused reagents and dispose of wastes as required by relevant regulations and codes. |  |  |  |  |  |
|  | Clean and store equipment. |  |  |  |  |  |
|  | Standardize Turbidity Meter by running known standards. |  |  |  |  |  |
|  | Analyzed Method blank after every 10 samples |  |  |  |  |  |
|  | Run the Laboratory Control samples of 10 FTU/NTU with every batch of 10 samples |  |  |  |  |  |
|  | check repeatability by reanalyzing a pre-analyzed sample once a day |  |  |  |  |  |
|  | Record Daily quality control data in QC Data sheet-I. |  |  |  |  |  |
|  | Note down the Results on analyst workbook. |  |  |  |  |  |
|  | Calculate uncertainty of theresults |  |  |  |  |  |
|  | Record the results on result record form . |  |  |  |  |  |
|  | Clear and restore work area. |  |  |  |  |  |
|  | Adopt precautionary measures for cuvette handling. |  |  |  |  |  |
|  | Follow Proper procedure for sample handling . |  |  |  |  |  |
|  | Check glassware cleanliness. Dirty glassware could also affect readings along with scratched or edged vials. |  |  |  |  |  |
| Competent |  |  | Not Yet Competent |  |  |  |

Knowledge Assessment

| Qualification | National Vocational Certificate-Level-3 in Water Quality & Resource Management (Water Quality Surveyor) |
| --- | --- |
| Competency Standard | Testing of Turbidity by Turbidity Meter |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: _________________ Candidate Signature: ___________ |
| Assessment Outcome | COMPETENT NOT YET COMPETENT Name of the Assessor: ________________________________________________ Assessor’s code: ____________________________________________________ Signature of the Assessor: _______________________ |

Candidate’s response is not required to be identical, but similar concepts and/or keywords must be used. Oral questioning may be used to clarify candidate understanding of topic and its application.

| Questions (Candidate confidently answered questions correctly and demonstrated understanding of the topics and their application) |  | Satisfactory | Not Satisfactory |
| --- | --- | --- | --- |
|  | What is meant by suspension? |  |  |
|  | What is the unit of Turbidity? |  |  |
|  | Can color of water sample effects turbidity? |  |  |
|  | What is the difference between turbidity and TSS? |  |  |
|  | What is the drinking water range of turbidity? |  |  |
|  | Which method is used for measuring turbidity? |  |  |
|  | What are Conventional Methods for measuring turbidity? |  |  |

Feedback to the Candidate

Candidate’s Signature__________________ Assessor’s Signature _________________