Instruction Sheet for the Candidate

| Qualification | National Vocational Certificate Level-4 in “Water Quality and Resource Management” (Lab Analyst) |
| --- | --- |
| Competency Standard | Develop Control Charts |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name______________________________________________________ Registration/Roll Number_______________________________________ |
| Guidance for Candidate | To meet this standard, you are required to complete the following within 04 Hrs. time frame (for practical demonstration & assessment): - Prepare Lab Control Sample for Control Chart - Analyse the Lab Control Sample - Set the Control Limits - Interpret the Control chart data |
| Time: 04 Hrs. | During a practical assessment, under observation by an assessor, you are required to |

| Minimum Evidence Required | Prepare Lab Control Sample for Control Chart - Considering the measurement range of various parameter/methods Prepare lab control samples by following steps - Prepare Stock solution from certified reference material/in-house standards. - Desired standard concentration is prepared from the stock solution by using the following formula C1V1 = C2V2 C1 = Initial Concentration V1 = Required Volume C2 = Concentration Required V2 = Final Volume Analyse the Lab Control Sample - Analyse the Lab control sample of each parameter 30 times by different analysts on different days. - Calculate the Mean as ∑ of results No. of Observations - calculate the Standard Deviation (SD) as Xi= No of Observations x¯=Mean Value N= No of readings Set the Control Limits Set the Control Limits as fallowing - Upper control limit UCL = x¯ + 3 s - Upper warning limit UWL = x + 2 s - Lower warning limit LWL = x - 2 s - Lower control limit LCL = x - 3 s Interpret the Control chart data - Control Limits: If one measurement exceeds the CL repeat the analyses immediately. If the repeat measurement is within the CL, continue the analyses. If it exceeds the CL, discontinue the analyses, identify the root cause and correct the problem. - Warning Limits: If two out of three successive points exceeds the WL, analyse another sample. If the next point is within the WL, continue analyses, if the next point exceeds the WL, evaluate potential bias and correct the problem. - Standard Deviation: If four out of five successive points exceeds 1s, or are in decreasing or in increasing order indicate the mean shift. Analyse another lab control sample for confirmation. If the next point is less than 1s, or changes the order, continue analyses, otherwise discontinue analyses and correct the problem. - Trending: If seven successive points are on the same side of the central line, indicate that there has been a shift in the mean; work can continue, however the cause of the shift needs to be identified and the limits re-set if justified |
| --- | --- |

Self-Assessment Checklist

| Candidate Name |  |
| --- | --- |
| Registration No. |  |
| Qualification | National Vocational Certificate Level-4 in “Water Quality and Resource Management” (Lab Analyst) |
| Competency Standard | Develop Control Charts |
| Purpose of Assessment | Formative Assessment |
| Assessment Task | - Prepare Lab Control Sample for Control Chart - Analyse the Lab Control Sample - Set the Control Limits - Interpret the Control chart data |

I can……………….

| Performance Criteria | Yes | No |
| --- | --- | --- |
| - Considering the measurement range of various parameter/methods Prepare lab control samples by following steps |  |  |
| - Prepare Stock solution from certified reference material/in-house standards. |  |  |
| - Analyse the Lab control sample of each parameter 30 times by different analysts on different days. |  |  |
| - If one measurement exceeds the CL repeat the analyses immediately. If the repeat measurement is within the CL, continue the analyses. If it exceeds the CL, discontinue the analyses, identify the root cause and correct the problem. |  |  |
| - If two out of three successive points exceeds the WL, analyse another sample. If the next point is within the WL, continue analyses, if the next point exceeds the WL, evaluate potential bias and correct the problem. |  |  |
| - If four out of five successive points exceeds 1s, or are in decreasing or in increasing order indicate the mean shift. Analyse another lab control sample for confirmation. If the next point is less than 1s, or changes the order, continue analyses, otherwise discontinue analyses and correct the problem. |  |  |
| - If seven successive points are on the same side of the central line, indicate that there has been a shift in the mean; work can continue, however the cause of the shift needs to be identified and the limits re-set if justified. |  |  |

Candidate’s Signature__________________ Assessor’s Signature____________________

Date: ______________________________

Assessors Judgment Guide

| Qualification | National Vocational Certificate Level-4 in “Water Quality and Resource Management” (Lab Analyst) |
| --- | --- |
| Competency Standard | Develop Control Charts |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: ______________________ Signature: _______________ |
| Assessment Outcome | COMPETENT NOT YET COMPETENT Name of the Assessor________________________________________________ Assessor’s code: ____________________________________________________ Signature: ________________________________ |

| Assessment Summary (to be filled by the assessor) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Activity | Method |  |  |  |  | Result |  |
| Nature of Activity | Written | Oral | Observation | Portfolio | Role Play | Competent | Not Yet Competent |
| Practical Skill Demonstration |  |  |  |  |  |  |  |
| Knowledge Assessment |  |  |  |  |  |  |  |
| Other Requirement |  |  |  |  |  |  |  |

Observation Checklist

| Assessment Task |  | - Prepare Lab Control Sample for Control Chart - Analyse the Lab Control Sample - Set the Control Limits - Interpret the Control chart data |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| During the practical assessment, candidate demonstrated the following: |  |  |  | Yes | No | Remarks |
|  | Considering the measurement range of various parameter/methods Prepare lab control samples by following steps |  |  |  |  |  |
|  | Prepare Stock solution from certified reference material/in-house standards. |  |  |  |  |  |
|  | Analyse the Lab control sample of each parameter 30 times by different analysts on different days. |  |  |  |  |  |
|  | If one measurement exceeds the CL repeat the analyses immediately. If the repeat measurement is within the CL, continue the analyses. If it exceeds the CL, discontinue the analyses, identify the root cause and correct the problem. |  |  |  |  |  |
|  | If two out of three successive points exceeds the WL, analyse another sample. If the next point is within the WL, continue analyses, if the next point exceeds the WL, evaluate potential bias and correct the problem. |  |  |  |  |  |
|  | If four out of five successive points exceeds 1s, or are in decreasing or in increasing order indicate the mean shift. Analyse another lab control sample for confirmation. If the next point is less than 1s, or changes the order, continue analyses, otherwise discontinue analyses and correct the problem. |  |  |  |  |  |
|  | If seven successive points are on the same side of the central line, indicate that there has been a shift in the mean; work can continue, however the cause of the shift needs to be identified and the limits re-set if justified. |  |  |  |  |  |
| Competent |  |  | Not Yet Competent |  |  |  |

Knowledge Assessment

| Qualification | National Vocational Certificate Level-4 in “Water Quality and Resource Management” (Lab Analyst) |
| --- | --- |
| Competency Standard | Develop Control Charts |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: _________________ Candidate Signature: ___________ |
| Assessment Outcome | COMPETENT NOT YET COMPETENT Name of the Assessor: ________________________________________________ Assessor’s code: ____________________________________________________ Signature of the Assessor: _______________________ |

Candidate’s response is not required to be identical, but similar concepts and/or keywords must be used. Oral questioning may be used to clarify candidate understanding of topic and its application.

| Questions (Candidate confidently answered questions correctly and demonstrated understanding of the topics and their application) |  | Satisfactory | Not Satisfactory |
| --- | --- | --- | --- |
|  | Write the formula for calculating the standard deviation? |  |  |
|  | Set the control limits for control charts based on standard deviation? |  |  |

Feedback to the Candidate

Candidate’s Signature__________________ Assessor’s Signature _________________