Instruction Sheet for the Candidate

| Qualification | National Vocational Certificate Level-4 in “Water Quality and Resource Management” (Lab Analyst) |
| --- | --- |
| Competency Standard | Quality Control Operation (Quality Control Technician) |
| Purpose of Assessment | Summative Assessment |
| Candidate Details | Name______________________________________________________ Registration/Roll Number_______________________________________ |
| Guidance for Candidate | To meet this standard, you are required to complete the following within 04 Hrs. time frame (for practical demonstration & assessment): Prepare, Analyze and evaluate the Audit Blink Sample - Prepare method blank from deionized water. - Analyze the method blank and record the result - Evaluate the method blank results Perform Calibration - Perform instrument calibration using standard solution - Perform initial calibration (IC) using standard solution or CRM further utilizing three desired standard solution for calibration range to generate standard curve - Perform Continued calibration verification (ccv) at specified intervals throughout the analytical sequence. Prepare and Analyze the Laboratory Control Sample - Prepare the laboratory control sample from solution of Certified Reference Material in de-ionized water as per standards. - Analyze the Lab Control in each batch after every 10 samples while plotting the results on control chart. - Evaluate the results Prepare and Analyze the Audit Blind Sample - Prepare standard solution for the desired parameter - Evaluate the results of Audit blind sample using formula. Perform anion cation balance - Perform the anion-cation balance by the following formula: % difference=100x∑Cations-∑Anions ∑Cations + ∑Anions The acceptance criteria is as fallows Anion Sum (meq/L) Acceptable difference 0-3.0 + 0.2meq/L 3.0-10.0 + 0.2% 10.0-800 + 5% Check Measured EC and ions Sums - Measure the electrical conductivity of the sample - sum the anions and cations - Both the anions and cations sum should be approximately 1/100 of the measured EC value - If either sum does not meet this criteria, re analyse the sample - The acceptance criteria is as fallows 100xanions(or cations) sum, meq/L= (0.9 – 1.1) EC Checked Measured TDS to EC ratio - Measure the electrical conductivity of the sample - Measure the TDS of the sample - Calculate the ratio of the TDS and EC as Measured TDS/conductivity= 0.55 – 0.7 - The measured TDS to EC ratio should be between 0.55 and 0.7. If it is outside these limits thee either measured TDS or measure conductivity is suspect; reanalyse the sample Prepare and Analyze the standard solution for Calculating Method Detection Limit - Prepare the standard solution four times the estimated limit - Analyze the standard solution 10 ten times - Calculate the MDL by using the following equation: Method Detection Limit (MDL) = t value x Std. deviation Where t = degree of freedom (n-1) Evaluate the MDL Results - If the MDL is much higher than the estimated detection limit, it may indicate problems with the procedure or with the analytical equipment, improperly made reagents and expiration dates for chemicals. Reprepare all the standards. Re analyse the standard and calculate the method detection limit. Prepare and Analyze Lab Control Sample for Control Chart - Prepare the standard solution for control chart - Analyze the lab control sample and calculate the mean and standard deviation Set the Control Limits - Set the control limit as per the standard deviation and mean. Interpret the Control chart data - Interpret the control chart data as per the rules set. |
| Time: 04 Hrs. | During a practical assessment, under observation by an assessor, you are required to |

Self-Assessment Checklist

| Candidate Name |  |
| --- | --- |
| Registration No. |  |
| Qualification | National Vocational Certificate Level-4 in “Water Quality and Resource Management” (Lab Analyst) |
| Competency Standard | Water and Waste Water Sampling and Field Monitoring |
| Purpose of Assessment | Formative Assessment |
| Assessment Task | - Prepare, Analyze and evaluate the Audit Blink Sample - Perform Calibration - Prepare and Analyze the Laboratory Control Sample - Prepare and Analyze the Audit Blind Sample - Perform anion cation balance - Check Measured EC and ions Sums - Checked Measured TDS to EC ratio - Prepare and Analyze the standard solution - Evaluate the MDL Results - Prepare and Analyze Lab Control Sample for Control Chart - Set the Control Limits - Interpret the Control chart data |

I can……………….

| Performance Criteria | Yes | No |
| --- | --- | --- |
| - Prepare method blank from deionized water. |  |  |
| - Analyze the method blank and record the result |  |  |
| - Evaluate the method blank results |  |  |
| - Perform instrument calibration using standard solution |  |  |
| - Perform initial calibration (IC) using standard solution or CRM further utilizing three desired standard solution for calibration range to generate standard curve |  |  |
| - Perform Continued calibration verification (ccv) at specified intervals throughout the analytical sequence. |  |  |
| - Prepare the laboratory control sample from solution of Certified Reference Material in de-ionized water as per standards. |  |  |
| - Analyze the Lab Control in each batch after every 10 samples while plotting the results on control chart. |  |  |
| - Evaluate the results |  |  |
| - Prepare standard solution for the desired parameter |  |  |
| - Evaluate the results of Audit blind sample using formula. |  |  |
| - Measure the electrical conductivity of the sample |  |  |
| - sum the anions and cations |  |  |
| - Both the anions and cations sum should be approximately 1/100 of the measured EC value |  |  |
| - If either sum does not meet this criteria, re analyse the sample |  |  |
| - Measure the electrical conductivity of the sample |  |  |
| - Measure the TDS of the sample |  |  |
| - Prepare the standard solution four times the estimated limit |  |  |
| - Analyze the standard solution 10 ten times |  |  |
| - Measure the electrical conductivity of the sample |  |  |
| - Prepare standard solution for the desired parameter |  |  |
| - Evaluate the results of Audit blind sample using formula. |  |  |
| - Measure the electrical conductivity of the sample |  |  |
| - sum the anions and cations |  |  |
| - Both the anions and cations sum should be approximately 1/100 of the measured EC value |  |  |
| - If either sum does not meet this criteria, re analyse the sample |  |  |
| - Measure the electrical conductivity of the sample |  |  |
| - Measure the TDS of the sample |  |  |
| - If the MDL is much higher than the estimated detection limit, it may indicate problems with the procedure or with the analytical equipment, improperly made reagents and expiration dates for chemicals. Reprepare all the standards. Re analyse the standard and calculate the method detection limit. |  |  |
| - Prepare the standard solution for control chart |  |  |
| - Analyze the lab control sample and calculate the mean and standard deviation |  |  |
| - Set the control limit as per the standard deviation and mean. |  |  |
| - Interpret the control char data as per the rules set. |  |  |

Candidate’s Signature__________________ Assessor’s Signature____________________

Date: _______________________________

Assessors Judgment Guide

| Qualification | National Vocational Certificate Level-4 in “Water Quality and Resource Management” (Lab Analyst) |
| --- | --- |
| Competency Standard | Water and Waste Water Sampling and Field Monitoring |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: ______________________ Signature: _______________ |
| Assessment Outcome | COMPETENT NOT YET COMPETENT Name of the Assessor________________________________________________ Assessor’s code: ____________________________________________________ Signature: ________________________________ |

| Assessment Summary (to be filled by the assessor) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Activity | Method |  |  |  |  | Result |  |
| Nature of Activity | Written | Oral | Observation | Portfolio | Role Play | Competent | Not Yet Competent |
| Practical Skill Demonstration |  |  |  |  |  |  |  |
| Knowledge Assessment |  |  |  |  |  |  |  |
| Other Requirement |  |  |  |  |  |  |  |

Observation Checklist

| Assessment Task |  | - Prepare sampling plan - Observe the site Details - Collect Sample form Ground water - Collect Sample for Quality Control - Perform on site Field Testing - Storage and Transportation of Samples |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| During the practical assessment, candidate demonstrated the following: |  |  |  | Yes | No | Remarks |
|  | Prepare method blank from deionized water. |  |  |  |  |  |
|  | Analyze the method blank and record the result |  |  |  |  |  |
|  | Evaluate the method blank results |  |  |  |  |  |
|  | Perform instrument calibration using standard solution |  |  |  |  |  |
|  | Perform initial calibration (IC) using standard solution or CRM further utilizing three desired standard solution for calibration range to generate standard curve |  |  |  |  |  |
|  | Perform Continued calibration verification (ccv) at specified intervals throughout the analytical sequence. |  |  |  |  |  |
|  | Prepare the laboratory control sample from solution of Certified Reference Material in de-ionized water as per standards. |  |  |  |  |  |
|  | Analyze the Lab Control in each batch after every 10 samples while plotting the results on control chart. |  |  |  |  |  |
|  | Evaluate the results |  |  |  |  |  |
|  | Prepare standard solution for the desired parameter |  |  |  |  |  |
|  | Evaluate the results of Audit blind sample using formula. |  |  |  |  |  |
|  | Measure the electrical conductivity of the sample |  |  |  |  |  |
|  | sum the anions and cations |  |  |  |  |  |
|  | Both the anions and cations sum should be approximately 1/100 of the measured EC value |  |  |  |  |  |
|  | If either sum does not meet this criteria, re analyse the sample |  |  |  |  |  |
|  | Measure the electrical conductivity of the sample |  |  |  |  |  |
|  | Measure the TDS of the sample |  |  |  |  |  |
|  | Prepare the standard solution four times the estimated limit |  |  |  |  |  |
|  | Analyze the standard solution 10 ten times |  |  |  |  |  |
|  | Measure the electrical conductivity of the sample |  |  |  |  |  |
|  | Prepare standard solution for the desired parameter |  |  |  |  |  |
|  | Evaluate the results of Audit blind sample using formula. |  |  |  |  |  |
|  | Measure the electrical conductivity of the sample |  |  |  |  |  |
|  | sum the anions and cations |  |  |  |  |  |
|  | Both the anions and cations sum should be approximately 1/100 of the measured EC value |  |  |  |  |  |
|  | If either sum does not meet this criteria, re analyse the sample |  |  |  |  |  |
|  | Measure the electrical conductivity of the sample |  |  |  |  |  |
|  | Measure the TDS of the sample |  |  |  |  |  |
|  | If the MDL is much higher than the estimated detection limit, it may indicate problems with the procedure or with the analytical equipment, improperly made reagents and expiration dates for chemicals. Reprepare all the standards. Re analyse the standard and calculate the method detection limit. |  |  |  |  |  |
|  | Prepare the standard solution for control chart |  |  |  |  |  |
|  | Analyze the lab control sample and calculate the mean and standard deviation |  |  |  |  |  |
| Competent |  |  | Not Yet Competent |  |  |  |

Knowledge Assessment

| Qualification | National Vocational Certificate Level-4 in “Water Quality and Resource Management” (Lab Analyst) |
| --- | --- |
| Competency Standard | Water and Waste Water Sampling and Field Monitoring |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: _________________ Candidate Signature: ___________ |
| Assessment Outcome | COMPETENT NOT YET COMPETENT Name of the Assessor: ________________________________________________ Assessor’s code: ____________________________________________________ Signature of the Assessor: _______________________ |

Candidate’s response is not required to be identical, but similar concepts and/or keywords must be used. Oral questioning may be used to clarify candidate understanding of topic and its application.

| Questions (Candidate confidently answered questions correctly and demonstrated understanding of the topics and their application) |  |  |  |  |  |  | Satisfactory |  |  | Not Satisfactory |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | What is the basis role of analyzing Method Blank? |  |  |  |  |  |  |  |
|  |  |  |  | Write the specific procedure for preparation of the method blank Sample? |  |  |  |  |  |  |  |
|  |  |  |  | Define Calibration and its purpose? |  |  |  |  |  |  |  |
|  |  |  |  | Name the two method of Calibration? |  |  |  |  |  |  |  |
|  |  |  | Write the basic principle of analyzing the lab control sample? |  |  |  |  |  |  |  |  |
|  |  |  | Describe the procedure for preparation of sample for Lab control? |  |  |  |  |  |  |  |  |
|  |  |  |  | What is the basic purpose of Audit blind Sample? |  |  |  |  |  |  |  |
|  |  |  |  | Write the method of evaluation for Audit blind Sample? |  |  |  |  |  |  |  |
|  |  |  |  | Calculate the anion cation balance through formula? |  |  |  |  |  |  |  |
|  |  |  |  | How to Calculate TDS of any Sample through formula? |  |  |  |  |  |  |  |
|  |  |  | What is the minimum concentration required for calculating MDL |  |  |  |  |  |  |  |  |
|  |  |  | Write the formula of MDL calculation? |  |  |  |  |  |  |  |  |
|  |  | Write the formula for calculating the standard deviation? |  |  |  |  |  |  |  |  |
|  |  | Set the control limits for control charts based on standard deviation? |  |  |  |  |  |  |  |  |

Feedback to the Candidate

Candidate’s Signature__________________ Assessor’s Signature _________________