Instruction Sheet for the Candidate

| Qualification | National Diploma in Water Quality & Resource Management-Level-5 |
| --- | --- |
| Competency Standard | Perform Testing/Determination of Sodium (Na) by Flame-photometric Method |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name______________________________________________________ Registration/Roll Number_______________________________________ |
| Guidance for Candidate | To meet this standard, you are required to complete the following within 04 Hrs. time frame (for practical demonstration & assessment): - Identify the appropriate test method, prepare samples and stock solution for testing - Prepare equipment/apparatus for testing - Perform tests on samples as per SOPs - Perform quality control checks - Record the results/ finalize the work - Adopt the precautions during testing work |
| Time: 04 Hrs. | During a practical assessment, under observation by an assessor, you are required to: |

| Minimum Evidence Required | Identify the appropriate test method, Prepare samples and stock solution for testing - Identify appropriate test method, Glassware equipment and safety requirements. - Record the chemicals and standards involved in analysis. - Maintain the temperature of sample and environmental condition. - Prepare 1000 ppm sodium stock solution from chemical salt.Prepare various standards of sodium by dilution method from stock solution. Prepare Equipment/Apparatus for testing - Set up test equipment Flam-photometer and/or reagents in accordance with the specified work instructions. - Conduct pre-use and safety checks for verification of method. - Check the calibration status of equipment. Perform calibration if required with prepared standards for sodium. Perform tests on samples as per SOPs - Turned on the fuel supply at the source and pressed power switch to on the meter. - The ignition switch is pressed to on the flam and adjusts the flam from rotating knob of fuel adjuster. - Select the sodium filter for determination of sodium. - The suction of photometer is checked by aspirating 10 ml deionized water per minute. - Standard solution is aspirated and adjusted to obtain the positive reading. - When blank and standard are set. Unknown samples aspirated and the results noted. Analyzed Three replicate for each sample. - Store unused reagents and dispose of wastes as required by relevant regulations and codes. Clean and store equipment. Perform Quality Control Checks - Analyze the Method blank sample. - Examine the Laboratory Control sample of known concentration. - Check repeatability by reanalyzing a pre-analyzed sample. - Record the quality control data. - Record the results/ Finalize the work - Note down the results on workbook. - Calculate uncertainty of the results and report the results with proper units. - Clear and restore work area. Adopt the precautions during testing work - Do not leave the instrument unattended while the flame is alight. - The top of instrument chimney unit becomes very hot when running don’t touch it. Never attempt to look down the chimney whilst the flame is running. Always use the inspection window. - If high concentration salt solution is aspirated for longer periods then deionized water aspirate prior to shut down. - Take care when preparing standards. Never use glass container to store calibration standards. - Always take sample from top half of the sample. The bottom half may contain sediments that could block fine tubing. |
| --- | --- |

Self-Assessment Checklist

| Candidate Name |  |
| --- | --- |
| Registration No. |  |
| Qualification | National Diploma in Water Quality & Resource Management-Level-5 |
| Competency Standard | Perform Testing/Determination of Sodium (Na) by Flame-photometric Method |
| Purpose of Assessment | Formative Assessment |
| Assessment Task | - Identify the appropriate test method, Prepare samples and stock solution for testing - Prepare Equipment/Apparatus for testing - Perform tests on samples as per SOPs - Perform Quality Control Checks - Record the results/ Finalize the work - Adopt the precautions during testing work |

I can……………….

| Performance Criteria | Yes | No |
| --- | --- | --- |
| - Identify appropriate test method, Glassware equipment and safety requirements. |  |  |
| - Record the chemicals and standards involved in analysis. |  |  |
| - Maintain the temperature of sample and environmental condition. |  |  |
| - Prepare 1000 ppm sodium stock solution from chemical salt. Prepare various standards of sodium by dilution method from stock solution. |  |  |
| - Set up test equipment Flam-photometer and/or reagents in accordance with the specified work instructions. |  |  |
| - Conduct pre-use and safety checks for verification of method |  |  |
| - Check the calibration status of equipment. Perform calibration if required with prepared standards for sodium. |  |  |
| - Turned on the fuel supply at the source and pressed power switch to on the meter. |  |  |
| - The ignition switch is pressed to on the Flam and adjusts the flam from rotating knob of fuel adjuster. |  |  |
| - Select the sodium filter for determination of sodium. |  |  |
| - The suction of photometer is checked by aspirating 10 ml deionized water per minute. |  |  |
| - Standard solution is aspirated and adjusted to obtain the positive reading. |  |  |
| - When blank and standard are set. Unknown samples aspirated and the results noted. Analyzed Three replicate for each sample. |  |  |
| - Store unused reagents and dispose of wastes as required by relevant regulations and codes. Clean and store equipment. |  |  |
| - Analyze the Method blank sample. |  |  |
| - Examine the Laboratory Control sample of known concentration. |  |  |
| - Check repeatability by reanalyzing a pre-analyzed sample. |  |  |
| - Record the quality control data. |  |  |
| - Note down the results on workbook. |  |  |
| - Calculate uncertainty of the results and report the results with proper units. |  |  |
| - Clear and restore work area. |  |  |
| - Do not leave the instrument unattended while the flame is alight. |  |  |
| - The top of instrument chimney unit becomes very hot when running don’t touch it. Never attempt to look down the chimney whilst the flame is running. Always use the inspection window. |  |  |
| - If high concentration salt solution is aspirated for longer periods then deionized water aspirate prior to shut down. |  |  |
| - Take care when preparing standards. Never use glass container to store calibration standards. |  |  |
| - Always take sample from top half of the sample. The bottom half may contain sediments that could block fine tubing. |  |  |

Candidate’s Signature__________________ Assessor’s Signature____________________

Date: _______________________________

Assessors Judgment Guide

| Qualification | National Diploma in Water Quality & Resource Management-Level-5 |
| --- | --- |
| Competency Standard | Perform Testing/Determination of Sodium (Na) by Flame-photometric Method |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: ______________________Signature: _______________ |
| Assessment Outcome | COMPETENT NOT YETCOMPETENT Name of the Assessor________________________________________________ Assessor’s code: ____________________________________________________ Signature: ________________________________ |

| Assessment Summary (to be filled by the assessor) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Activity | Method |  |  |  |  | Result |  |
| Nature of Activity | Written | Oral | Observation | Portfolio | Role Play | Competent | Not Yet Competent |
| Practical Skill Demonstration |  |  |  |  |  |  |  |
| Knowledge Assessment |  |  |  |  |  |  |  |
| Other Requirement |  |  |  |  |  |  |  |

Observation Checklist

| Assessment Task |  | - Identify the appropriate test method, Prepare samples and stock solution for testing - Prepare Equipment/Apparatus for testing - Perform tests on samples as per SOPs - Perform Quality Control Checks - Record the results/ Finalize the work - Adopt the precautions during testing work |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| During the practical assessment, candidate demonstrated the following: |  |  |  | Yes | No | Remarks |
| - Identify appropriate test method, Glassware equipment and safety requirements. |  |  |  |  |  |
| - Record the chemicals and standards involved in analysis. |  |  |  |  |  |
| - Maintain the temperature of sample and environmental condition. |  |  |  |  |  |
| - Prepare 1000 ppm sodium stock solution from chemical salt. Prepare various standards of sodium by dilution method from stock solution. |  |  |  |  |  |
| - Set up test equipment Flam-photometer and/or reagents in accordance with the specified work instructions. |  |  |  |  |  |
| - Conduct pre-use and safety checks for verification of method |  |  |  |  |  |
| - Check the calibration status of equipment. Perform calibration if required with prepared standards for sodium. |  |  |  |  |  |
| - Turned on the fuel supply at the source and pressed power switch to on the meter. |  |  |  |  |  |
| - The ignition switch is pressed to on the flam and adjusts the flam from rotating knob of fuel adjuster. |  |  |  |  |  |
| - Select the sodium filter for determination of sodium. |  |  |  |  |  |
| - The suction of photometer is checked by aspirating 10 ml deionized water per minute. |  |  |  |  |  |
| - Standard solution is aspirated and adjusted to obtain the positive reading. |  |  |  |  |  |
| - When blank and standard are set. Unknown samples aspirated and the results noted. Analyzed Three replicate for each sample. |  |  |  |  |  |
| - Store unused reagents and dispose of wastes as required by relevant regulations and codes. Clean and store equipment. |  |  |  |  |  |
| - Analyze the Method blank sample. |  |  |  |  |  |
| - Examine the Laboratory Control sample of known concentration. |  |  |  |  |  |
| - Check repeatability by reanalyzing a pre-analyzed sample. |  |  |  |  |  |
| - Record the quality control data. |  |  |  |  |  |
| - Note down the results on workbook. |  |  |  |  |  |
| - Calculate uncertainty of the results and report the results with proper units. |  |  |  |  |  |
| - Clear and restore work area. |  |  |  |  |  |
| - Do not leave the instrument unattended while the flame is alight. |  |  |  |  |  |
| - The top of instrument chimney unit becomes very hot when running don’t touch it. Never attempt to look down the chimney whilst the flame is running. Always use the inspection window. |  |  |  |  |  |
| - If high concentration salt solution is aspirated for longer periods then deionized water aspirate prior to shut down. |  |  |  |  |  |
| - Take care when preparing standards. Never use glass container to store calibration standards. |  |  |  |  |  |
| - Always take sample from top half of the sample. The bottom half may contain sediments that could block fine tubing. |  |  |  |  |  |
| Competent |  |  | Not Yet Competent |  |  |  |

Knowledge Assessment

| Qualification | National Diploma in Water Quality & Resource Management-Level-5 |
| --- | --- |
| Competency Standard | Perform Testing/Determination of Sodium (Na) by Flame-photometric Method |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: _________________ Candidate Signature: ___________ |
| Assessment Outcome | COMPETENT NOT YETCOMPETENT Name of the Assessor: ________________________________________________ Assessor’s code: ____________________________________________________ Signature of the Assessor: _______________________ |

Candidate’s response is not required to be identical, but similar concepts and/or keywords must be used. Oral questioning may be used to clarify candidate understanding of topic and its application.

| Questions (Candidate confidently answered questions correctly and demonstrated understanding of the topics and their application) |  | Satisfactory | Not Satisfactory |
| --- | --- | --- | --- |
|  | What is the basic principle of Flam-photometric sodium measurement |  |  |
|  | What color flame is seen when sodium ignited? |  |  |
|  | What is the function of fuel adjuster flame photometer |  |  |
|  | Given below is the diagram of Flame emission photometers. Identify the unmarked component |  |  |
|  | What are the drinking water standards for sodium? |  |  |

Feedback to the Candidate

Candidate’s Signature__________________ Assessor’s Signature _________________