Instruction Sheet for the Candidate

| Qualification | National Diploma-Level-5 in “Water Quality and Resource Management” (Technician-Supervisor) |
| --- | --- |
| Competency Standard | Perform Testing/Determination of Copper (Cu) by Atomic Absorption Spectrometer |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name______________________________________________________ Registration/Roll Number_______________________________________ |
| Guidance for Candidate | To meet this standard, you are required to complete the following within 04 Hrs. time frame (for practical demonstration & assessment): - Identify the appropriate test method, Prepare samples and stock solution for testing - Prepare Equipment/Apparatus for testing - Perform tests on samples as per SOPs - Perform Quality Control Checks - Record the results/ Finalize the work - Adopt the precautions during testing work |
| Time: 04 Hrs. | During a practical assessment, under observation by an assessor, you are required to |

| Minimum Evidence Required | Identify the appropriate test method, Prepare samples and stock solution for testing - Identify appropriate test method, Glassware equipment and safety requirements - Check the sample label for the requirement of analysis of Copper. - Keeps the sample at room temperature for few minutes. - Check for the availability of Copper Standard solution 0.2-1 ppm otherwise prepare as per standard procedure from reference stock solution of 1000 ppm Prepare Equipment/Apparatus for testing - Set up test equipment AAS and/or reagents in accordance with the specified work instructions. - Conduct pre-use and safety checks. - Check the calibration status of equipment. Perform calibration if required. - Conduct pre-use and safety checks. Perform tests on samples as per SOPs - Open the AAS Software select Flam Mode. Load Method for Copper or draw new calibration curve. - Adjust the parameter (energy of the lamps, auto sampler parameters and adjustment of air acetylene flam parameter according to the manufacturer instructions). - Insert the sample codes in sample table along with giving the sample cup positions as in autosampler. - Take sample in sample cup of autosampler up to the mark and put in specific position of sample tray of auto sampler. - Ignite the flam by clicking on the ignition option in main option “Flam”. - Run Sample from software or click on option “Start Conc.” after selecting the range of samples in table by giving working area in case of autosampler used. - Feed each sample one by one manually in case of working without autosampler - Analyze three replicate for each sample. - Store unused reagents and dispose of wastes as required by relevant regulations and codes. - Clean and store equipment. Perform Quality Control Checks - Analyze Blank and Certified Reference Standard after every ten samples. - Analyze pre-analyzed sample with every analytical batch. - Analyze Spiked Matrix sample after every two months - Analyze the Control Sample and display data is regularly on control charts. - Record the daily quality control data in QC Data sheet. - Record the results/ Finalize the work - Note down the Results on analyst workbook. - Calculate uncertainty of the results - Record the results on result record form and submit to reporting section - Clear and restore work area. Adopt the precautions during testing work - All the glassware should be washed with 1N HNO3 before preparation of standard and samples. - Auto-pipettes should be used carefully to take accurate volume during the analysis. - Deionized water used should be of high quality having EC <1.5µS/cm. |
| --- | --- |

Self-Assessment Checklist

| Candidate Name |  |
| --- | --- |
| Registration No. |  |
| Qualification | National Diploma-Level-5 in “Water Quality and Resource Management” (Technician-Supervisor) |
| Competency Standard | Perform Testing/Determination of Copper (Cu) by Atomic Absorption Spectrometer |
| Purpose of Assessment | Formative Assessment |
| Assessment Task | - Identify the appropriate test method, Prepare samples and stock solution for testing - Prepare Equipment/Apparatus for testing - Perform tests on samples as per SOPs - Perform Quality Control Checks - Record the results/ Finalize the work - Adopt the precautions during testing work |

I can……………….

| Performance Criteria | Yes | No |
| --- | --- | --- |
| - Identify appropriate test method, Glassware equipment and safety requirements |  |  |
| - Check the sample label for the requirement of analysis of Copper. |  |  |
| - Keeps the sample at room temperature for few minutes. |  |  |
| - Check for the availability of Copper Standard solution 0.2-1 ppm otherwise prepare as per standard procedure from reference stock solution of 1000 ppm |  |  |
| - Set up test equipment AAS and/or reagents in accordance with the specified work instructions. |  |  |
| - Conduct pre-use and safety checks. |  |  |
| - Check the calibration status of equipment. Perform calibration if required. |  |  |
| - Conduct pre-use and safety checks. |  |  |
| - Open the AAS Software select Flam Mode. Load Method for Copper or draw new calibration curve. |  |  |
| - Adjust the parameter (energy of the lamps, auto sampler parameters and adjustment of air acetylene flam parameter according to the manufacturer instructions). |  |  |
| - Insert the sample codes in sample table along with giving the sample cup positions as in autosampler. |  |  |
| - Take sample in sample cup of autosampler up to the mark and put in specific position of sample tray of auto sampler. |  |  |
| - Ignite the flam by clicking on the ignition option in main option “Flam”. |  |  |
| - Run Sample from software or click on option “Start Conc.” after selecting the range of samples in table by giving working area in case of autosampler used. |  |  |
| - Feed each sample one by one manually in case of working without autosampler |  |  |
| - Analyze three replicate for each sample. |  |  |
| - Store unused reagents and dispose of wastes as required by relevant regulations and codes. |  |  |
| - Clean and store equipment. |  |  |
| - Clean and store equipment. |  |  |
| - Analyze Blank and Certified Reference Standard after every ten samples. |  |  |
| - Analyze pre-analyzed sample with every analytical batch. |  |  |
| - Analyze Spiked Matrix sample after every two months |  |  |
| - Analyze the Control Sample and display data is regularly on control charts. |  |  |
| - Record the daily quality control data in QC Data sheet. |  |  |
| - Note down the Results on analyst workbook. |  |  |
| - Calculate uncertainty of the results |  |  |
| - Record the results on result record form and submit to reporting section |  |  |
| - Clear and restore work area. |  |  |
| - All the glassware should be washed with 1N HNO3 before preparation of standard and samples. |  |  |
| - Auto-pipettes should be used carefully to take accurate volume during the analysis. |  |  |
| - Deionized water used should be of high quality having EC <1.5µS/cm. |  |  |

Candidate’s Signature__________________ Assessor’s Signature____________________

Date: _______________________________

Assessors Judgment Guide

| Qualification | National Diploma-Level-5 in “Water Quality and Resource Management” (Technician-Supervisor) |
| --- | --- |
| Competency Standard | Perform Testing/Determination of Copper (Cu) by Atomic Absorption Spectrometer |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: ______________________Signature: _______________ |
| Assessment Outcome | COMPETENT NOT YETCOMPETENT Name of the Assessor________________________________________________ Assessor’s code: ____________________________________________________ Signature: ________________________________ |

| Assessment Summary (to be filled by the assessor) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Activity | Method |  |  |  |  | Result |  |
| Nature of Activity | Written | Oral | Observation | Portfolio | Role Play | Competent | Not Yet Competent |
| Practical Skill Demonstration |  |  |  |  |  |  |  |
| Knowledge Assessment |  |  |  |  |  |  |  |
| Other Requirement |  |  |  |  |  |  |  |

Observation Checklist

| Assessment Task |  | - Identify the appropriate test method, Prepare samples and stock solution for testing - Prepare Equipment/Apparatus for testing - Perform tests on samples as per SOPs - Perform Quality Control Checks - Record the results/ Finalize the work - Adopt the precautions during testing work |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| During the practical assessment, candidate demonstrated the following: |  |  |  | Yes | No | Remarks |
|  | Identify appropriate test method, Glassware equipment and safety requirements |  |  |  |  |  |
|  | Check the sample label for the requirement of analysis of Copper. |  |  |  |  |  |
|  | Keeps the sample at room temperature for few minutes. |  |  |  |  |  |
|  | Check for the availability of Copper Standard solution 0.2-1 ppm otherwise prepare as per standard procedure from reference stock solution of 1000 ppm |  |  |  |  |  |
|  | Set up test equipment AAS and/or reagents in accordance with the specified work instructions. |  |  |  |  |  |
|  | Conduct pre-use and safety checks. |  |  |  |  |  |
|  | Check the calibration status of equipment. Perform calibration if required. |  |  |  |  |  |
|  | Conduct pre-use and safety checks. |  |  |  |  |  |
|  | Open the AAS Software select Flam Mode. Load Method for Copper or draw new calibration curve. |  |  |  |  |  |
|  | Adjust the parameter (energy of the lamps, auto sampler parameters and adjustment of air acetylene flam parameter according to the manufacturer instructions). |  |  |  |  |  |
|  | Insert the sample codes in sample table along with giving the sample cup positions as in autosampler. |  |  |  |  |  |
|  | Take sample in sample cup of autosampler up to the mark and put in specific position of sample tray of auto sampler. |  |  |  |  |  |
|  | Ignite the flam by clicking on the ignition option in main option “Flam”. |  |  |  |  |  |
|  | Run Sample from software or click on option “Start Conc.” after selecting the range of samples in table by giving working area in case of autosampler used. |  |  |  |  |  |
|  | Feed each sample one by one manually in case of working without autosampler |  |  |  |  |  |
|  | Analyze three replicate for each sample. |  |  |  |  |  |
|  | Store unused reagents and dispose of wastes as required by relevant regulations and codes. |  |  |  |  |  |
|  | Clean and store equipment. |  |  |  |  |  |
|  | Clean and store equipment. |  |  |  |  |  |
|  | Analyze Blank and Certified Reference Standard after every ten samples. |  |  |  |  |  |
|  | Analyze pre-analyzed sample with every analytical batch. |  |  |  |  |  |
|  | Analyze Spiked Matrix sample after every two months |  |  |  |  |  |
|  | Analyze the Control Sample and display data is regularly on control charts. |  |  |  |  |  |
|  | Record the daily quality control data in QC Data sheet. |  |  |  |  |  |
|  | Note down the Results on analyst workbook. |  |  |  |  |  |
|  | Calculate uncertainty of the results |  |  |  |  |  |
|  | Record the results on result record form and submit to reporting section |  |  |  |  |  |
|  | Clear and restore work area. |  |  |  |  |  |
|  | All the glassware should be washed with 1N HNO3 before preparation of standard and samples. |  |  |  |  |  |
|  | Auto-pipettes should be used carefully to take accurate volume during the analysis. |  |  |  |  |  |
|  | Deionized water used should be of high quality having EC <1.5µS/cm. |  |  |  |  |  |
| Competent |  |  | Not Yet Competent |  |  |  |

Knowledge Assessment

| Qualification | National Diploma-Level-5 in “Water Quality and Resource Management” (Technician-Supervisor) |
| --- | --- |
| Competency Standard | Perform Testing/Determination of Copper (Cu) by Atomic Absorption Spectrometer |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: _________________ Candidate Signature: ___________ |
| Assessment Outcome | COMPETENT NOT YETCOMPETENT Name of the Assessor: ________________________________________________ Assessor’s code: ____________________________________________________ Signature of the Assessor: _______________________ |

Candidate’s response is not required to be identical, but similar concepts and/or keywords must be used. Oral questioning may be used to clarify candidate understanding of topic and its application.

| Questions (Candidate confidently answered questions correctly and demonstrated understanding of the topics and their application) |  | Satisfactory | Not Satisfactory |
| --- | --- | --- | --- |
|  | Why Copper categorized as heavy metal |  |  |
|  | What are the potential sources of Copper in Water |  |  |
|  | Mention different techniques being used for the analysis of Copper in water |  |  |
|  | What are the possible health implications of Copper if high in drinking water |  |  |
|  | How we can remove the Copper from contaminated drinking and wastewater |  |  |
| 7. | At which wavelength copper is being analyzed on Atomic Absorption Spectrometer |  |  |
|  | What is the basic principle of copper analysis on flame mode of AAS |  |  |

Feedback to the Candidate

Candidate’s Signature__________________ Assessor’s Signature _________________