Instruction Sheet for the Candidate

| Qualification | National Diploma- Level-5 in “Water Quality and Resource Management” (Technician-Supervisor) |
| --- | --- |
| Competency Standard | Perform Testing/Determination of Biochemical Oxygen Demand (BOD) by Monometric Pressure Measurement Method |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name______________________________________________________ Registration/Roll Number_______________________________________ |
| Guidance for Candidate | To meet this standard, you are required to complete the following within 04 Hrs. Time frame (for practical demonstration & assessment): - Identify the appropriate test method, Prepare Standard required for testing - Prepare Equipment/Apparatus for testing - Perform tests on samples as per SOPs - Perform Quality Control Checks - Record the results/ Finalize the work - Adopt the precautions during testing work |
| Time: 04 Hrs. | During a practical assessment, under observation by an assessor, you are required to |

| Minimum Evidence Required | Identify the appropriate test method, Prepare samples and stock solution for testing - Identify appropriate test method, Glassware equipment and safety requirements. - Maintain the temperature of sample and working environment before analysis. - Maintain the optimum pH value for BOD measurement is between 6.5 and 7.5.The high or lower pH value can be readjusted by adding (1M) HCl or NaOH. - Homogenize the sample or filtered before analysis. - Perform nitrification inhibition if required. - Clean BOD bottles and stirring rods with teepol (detergent) and rinsed with hot water and finally rinsed with deionized water. - Identify appropriate test method, Glassware, Equipment and safety requirements.. Prepare Equipment/Apparatus for testing - Set up test equipment BOD Incubator and/or reagents in accordance with the specified work instructions. - Conduct pre-use and safety checks. - Incubate the standard with every batch of sample. Perform tests on samples as per SOPs - Set the Temperature of the thermostat cabinet at 20oC. - Carry out the Pre-treatment of the sample if required (e.g. Setting of pH, Filtering). - Select the measurement range for the sample under test and take the appropriate volume of sample as per manual instruction. - Measure the volume of the sample precisely by using the overflow measurement flask and transferred to BOD bottle. If necessary, add nitrification inhibitor or add seed in order to provide microbial growth if required. - Insert magnetic stirring rods in the BOD bottle and place drops of KOH solution into the seal gasket as per standard. Then insert the gasket in the rack of the bottle. - Incubate the sample and standard by selecting the required range by from instrument program 5 days - Store unused reagents and dispose of wastes as required by relevant regulations and codes. Perform Quality Control Checks - Analyze CRMs in order to check the calibration of the measuring system with every badge of samples. - Check the quality of dilution water by incubation. The DO uptake in 5 days at 20oC should not be more than (0.2 mg/l) preferably not more than 0.1/lit). - Measure the BOD of the seeding material and dilution water to check the quality of seed and dilution water. - Record the results/ Finalize the work - Record and report the Results with proper units and uncertainty. Adopt the precautions during testing work - The sample must never come in contact with the potassium Hydroxide solution. - Never use grease or any other lubricant as an additional sealing agent, for the BOD sensors or for the seal gasket. - Tighten BOD sensors carefully otherwise it will not detect the generated pressure. - Before measurement begins, the prepared sample must be brought to desired temperature (20oC). - Once the measurement has started never disturb the position of bottles otherwise you will lose the stored data. |
| --- | --- |

Self-Assessment Checklist

| Candidate Name |  |
| --- | --- |
| Registration No. |  |
| Qualification | National Diploma- Level-5 in “Water Quality and Resource Management” (Technician-Supervisor) |
| Competency Standard | Perform Testing/Determination of Biochemical Oxygen Demand (BOD) by Manometric Pressure Measurement Method |
| Purpose of Assessment | Formative Assessment |
| Assessment Task | - Identify the appropriate test method, Prepare samples and stock solution for testing - Prepare Equipment/Apparatus for testing - Perform tests on samples as per SOPs - Perform Quality Control Checks - Record the results/ Finalize the work - Adopt the precautions during testing work |

I can……………….

| Performance Criteria | Yes | No |
| --- | --- | --- |
| - Identify appropriate test method, Glassware equipment and safety requirements. |  |  |
| - Maintain the temperature of sample and working environment before analysis. |  |  |
| - Maintain the optimum pH value for BOD measurement is between 6.5 and 7.5.The high or lower pH value can be readjusted by adding (1M) HCl or NaOH. |  |  |
| - Homogenize the sample or filtered before analysis. |  |  |
| - Perform nitrification inhibition if required. |  |  |
| - Clean BOD bottles and stirring rods with teepol (detergent) and rinsed with hot water and finally rinsed with deionized water. |  |  |
| - Identify appropriate test method, Glassware, Equipment and safety requirements.. |  |  |
| - Set up test equipment BOD Incubator and/or reagents in accordance with the specified work instructions. |  |  |
| - Conduct pre-use and safety checks. |  |  |
| - Incubate the standard with every batch of sample. |  |  |
| - Set the Temperature of the thermostat cabinet at 20oC. |  |  |
| - Carry out the Pre-treatment of the sample if required (e.g. Setting of pH, Filtering). |  |  |
| - Select the measurement range for the sample under test and take the appropriate volume of sample as per manual instruction. |  |  |
| - Measure the volume of the sample precisely by using the overflow measurement flask and transferred to BOD bottle. If necessary, add nitrification inhibitor or add seed in order to provide microbial growth if required. |  |  |
| - Insert magnetic stirring rods in the BOD bottle and place drops of KOH solution into the seal gasket as per standard. Then insert the gasket in the rack of the bottle. |  |  |
| - Incubate the sample and standard by selecting the required range by from instrument program 5 days |  |  |
| - Store unused reagents and dispose of wastes as required by relevant regulations and codes. |  |  |
| - Analyze CRMs in order to check the calibration of the measuring system with every badge of samples. |  |  |
| - Check the quality of dilution water by incubation. The DO uptake in 5 days at 20oC should not be more than (0.2 mg/l) preferably not more than 0.1/lit). |  |  |
| - Measure the BOD of the seeding material and dilution water to check the quality of seed and dilution water. |  |  |
| - Record and report the Results with proper units and uncertainty. |  |  |
| - The sample must never come in contact with the potassium Hydroxide solution. |  |  |
| - Never use grease or any other lubricant as an additional sealing agent, for the BOD sensors or for the seal gasket. |  |  |
| - Tighten BOD sensors carefully otherwise it will not detect the generated pressure. |  |  |
| - Before measurement begins, the prepared sample must be brought to desired temperature (20oC). |  |  |
| - Once the measurement has started never disturb the position of bottles otherwise you will lose the stored data. |  |  |

Candidate’s Signature__________________ Assessor’s Signature____________________

Date: _______________________________

Assessors Judgment Guide

| Qualification | National Diploma- Level-5 in “Water Quality and Resource Management” (Technician-Supervisor) |
| --- | --- |
| Competency Standard | Perform Testing/Determination of Biochemical Oxygen Demand (BOD) by Manometric Pressure Measurement Method |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: ______________________Signature: _______________ |
| Assessment Outcome | COMPETENT NOT YETCOMPETENT Name of the Assessor________________________________________________ Assessor’s code: ____________________________________________________ Signature: ________________________________ |

| Assessment Summary (to be filled by the assessor) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Activity | Method |  |  |  |  | Result |  |
| Nature of Activity | Written | Oral | Observation | Portfolio | Role Play | Competent | Not Yet Competent |
| Practical Skill Demonstration |  |  |  |  |  |  |  |
| Knowledge Assessment |  |  |  |  |  |  |  |
| Other Requirement |  |  |  |  |  |  |  |

Observation Checklist

| Assessment Task |  | - Identify the appropriate test method, Prepare samples and stock solution for testing - Prepare Equipment/Apparatus for testing - Perform tests on samples as per SOPs - Perform Quality Control Checks - Record the results/ Finalize the work - Adopt the precautions during testing work |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| During the practical assessment, candidate demonstrated the following: |  |  |  | Yes | No | Remarks |
|  | Identify appropriate test method, Glassware equipment and safety requirements. |  |  |  |  |  |
|  | Maintain the temperature of sample and working environment before analysis. |  |  |  |  |  |
|  | Maintain the optimum pH value for BOD measurement is between 6.5 and 7.5.The high or lower pH value can be readjusted by adding (1M) HCl or NaOH. |  |  |  |  |  |
|  | Homogenize the sample or filtered before analysis. |  |  |  |  |  |
|  | Perform nitrification inhibition if required. |  |  |  |  |  |
|  | Clean BOD bottles and stirring rods with teepol (detergent) and rinsed with hot water and finally rinsed with deionized water. |  |  |  |  |  |
|  | Identify appropriate test method, Glassware, Equipment and safety requirements.. |  |  |  |  |  |
|  | Set up test equipment BOD Incubator and/or reagents in accordance with the specified work instructions. |  |  |  |  |  |
|  | Conduct pre-use and safety checks. |  |  |  |  |  |
|  | Incubate the standard with every batch of sample. |  |  |  |  |  |
|  | Set the Temperature of the thermostat cabinet at 20oC. |  |  |  |  |  |
|  | Carry out the Pre-treatment of the sample if required (e.g. Setting of pH, Filtering). |  |  |  |  |  |
|  | Select the measurement range for the sample under test and take the appropriate volume of sample as per manual instruction. |  |  |  |  |  |
|  | Measure the volume of the sample precisely by using the overflow measurement flask and transferred to BOD bottle. If necessary, add nitrification inhibitor or add seed in order to provide microbial growth if required. |  |  |  |  |  |
|  | Insert magnetic stirring rods in the BOD bottle and place drops of KOH solution into the seal gasket as per standard. Then insert the gasket in the rack of the bottle. |  |  |  |  |  |
|  | Incubate the sample and standard by selecting the required range by from instrument program 5 days |  |  |  |  |  |
|  | Store unused reagents and dispose of wastes as required by relevant regulations and codes. |  |  |  |  |  |
|  | Analyze CRMs in order to check the calibration of the measuring system with every badge of samples. |  |  |  |  |  |
|  | Check the quality of dilution water by incubation. The DO uptake in 5 days at 20oC should not be more than (0.2 mg/l) preferably not more than 0.1/lit). |  |  |  |  |  |
|  | Measure the BOD of the seeding material and dilution water to check the quality of seed and dilution water. |  |  |  |  |  |
|  | Record and report the Results with proper units and uncertainty. |  |  |  |  |  |
|  | The sample must never come in contact with the potassium Hydroxide solution. |  |  |  |  |  |
|  | Never use grease or any other lubricant as an additional sealing agent, for the BOD sensors or for the seal gasket. |  |  |  |  |  |
|  | Tighten BOD sensors carefully otherwise it will not detect the generated pressure. |  |  |  |  |  |
|  | Before measurement begins, the prepared sample must be brought to desired temperature (20oC). |  |  |  |  |  |
|  | Once the measurement has started never disturb the position of bottles otherwise you will lose the stored data. |  |  |  |  |  |
| Competent |  |  | Not Yet Competent |  |  |  |

Knowledge Assessment

| Qualification | National Diploma- Level-5 in “Water Quality and Resource Management” (Technician-Supervisor) |
| --- | --- |
| Competency Standard | Perform Testing/Determination of Biochemical Oxygen Demand (BOD) by Monometric Pressure Measurement Method |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: _________________ Candidate Signature: ___________ |
| Assessment Outcome | COMPETENT NOT YETCOMPETENT Name of the Assessor: ________________________________________________ Assessor’s code: ____________________________________________________ Signature of the Assessor: _______________________ |

Candidate’s response is not required to be identical, but similar concepts and/or keywords must be used. Oral questioning may be used to clarify candidate understanding of topic and its application.

| Questions (Candidate confidently answered questions correctly and demonstrated understanding of the topics and their application) |  | Satisfactory | Not Satisfactory |
| --- | --- | --- | --- |
|  | Define Waste water |  |  |
|  | What is BOD |  |  |
|  | Explain the rule of KOH during sample incubation |  |  |
|  | Write down the volume required for range of 0-200 |  |  |
|  | Define seeding |  |  |
|  | Why sample is incubated in amber colored bottle? |  |  |
|  | How will you maintain the pH of sample in the specific range? |  |  |
| 777 |  |  |  |

Feedback to the Candidate

Candidate’s Signature__________________ Assessor’s Signature _________________