Instruction Sheet for the Candidate

| Qualification | National Diploma- Level-5 in “Water Quality and Resource Management” (Technician-Supervisor) |
| --- | --- |
| Competency Standard | Perform Testing/Determination of Oil and Grease in Wastewater |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name______________________________________________________ Registration/Roll Number_______________________________________ |
| Guidance for Candidate | To meet this standard, you are required to complete the following within 04 Hrs. Time frame (for practical demonstration & assessment): - Identify the appropriate test method, Prepare Standard required for testing - Prepare Equipment/Apparatus for testing - Perform tests on samples as per SOPs - Perform Quality Control Checks - Record the results/ Finalize the work - Adopt the precautions during testing work |
| Time: 04 Hrs. | During a practical assessment, under observation by an assessor, you are required to |

| Minimum Evidence Required | Identify the appropriate test method, Prepare samples and stock solution for testing - Identify appropriate test method, Glassware equipment and safety requirements. - Maintain the temperature of sample and working environment before analysis. - Maintain the pH of sample. - Maintain the pH to <=2 by adding Hydrochloric acid - Shake the glass sample bottle to assure good mixing of sample - Prepare Zero liquid, span liquid and OCB Standard for Calibration, Pure Solvent S-315 is used as zero solvent. Prepare Equipment/Apparatus for testing - Set up test equipment OCMA-350 and reagents in accordance with the specified work instructions. - Perform calibration, zero, span and measurement of zero solvent in terms of mg/L. - Select specific units on OCMA-350 for oil in water analysis Perform tests on samples as per SOPs - Accurately pipette sample and Solvent S-316 in 40ml vial - Select the measurement range for the sample under test and take the appropriate volume of sample as per manual instruction. - Perform sample extraction with vigorous shaking for break between the solvent on bottom and water on the top. - In case of emulsion treat the extract to break it. Carefully pipette the as much solvent/extract as possible without taking the water layer in a beaker. - Break the emulsion by adding specific chemical Sodium Sulfate by stirring with glass rod - Put a conditioned 11cm No. 40 whatman filter paper into a glass funnel and filter the extract to remove the Na2SO4 - If less than 6ml of solvent/extract are recovered, add a measured amount of solvent, and account for dilution in the results. - After pipetting 6ml of solvent/extract fill and place the OCMA-350 cell into equipment for measurement. - For measurement of Total Petroleum Hydrocarbon (TPH). Treat the extract with conditioned Silica Gel, filtered with No. 40 Whatman filter paper and premeasure. - The difference, Oil and Grease minus Total Petroleum Hydrocarbon, is the non-petroleum Hydrocarbon. - Remove the cell from OCMA-350 and empty for next measurement. Perform Quality Control Checks - Perform the zero and span calibration before the measurement. Record the results/ Finalize the work - Record and report the Results with proper units and uncertainty. Adopt the precautions during testing work - Close the Cover on the unit to avoid dust. - When filling the cell, make sure there are no bubbles in the cell. - Do not touch the ends of a cell. - Do not spill the solvent on the measuring unit during cell insertion. - For accurate measuring results, calibration and measurement should be started after the “warmup” lamp off. - Wear gloves when handling it and avoid inhaling of vapors or ingest liquid. |
| --- | --- |

Self-Assessment Checklist

| Candidate Name |  |
| --- | --- |
| Registration No. |  |
| Qualification | National Diploma- Level-5 in “Water Quality and Resource Management” (Technician-Supervisor) |
| Competency Standard | Perform Testing/Determination of Oil and Grease in Wastewater by |
| Purpose of Assessment | Formative Assessment |
| Assessment Task | - Identify the appropriate test method, Prepare samples and stock solution for testing - Prepare Equipment/Apparatus for testing - Perform tests on samples as per SOPs - Perform Quality Control Checks - Record the results/ Finalize the work - Adopt the precautions during testing work |

I can……………….

| Performance Criteria | Yes | No |
| --- | --- | --- |
| - Identify appropriate test method, Glassware equipment and safety requirements. |  |  |
| - Maintain the temperature of sample and working environment before analysis. |  |  |
| - Maintain the pH of sample. |  |  |
| - Maintain the pH to <=2 by adding Hydrochloric acid |  |  |
| - Shake the glass sample bottle to assure good mixing of sample |  |  |
| - Prepare Zero liquid, span liquid and OCB Standard for Calibration, Pure Solvent S-315 is used as zero solvent. |  |  |
| - Set up test equipment OCMA-350 and reagents in accordance with the specified work instructions. |  |  |
| - Perform calibration, zero, span and measurement of zero solvent in terms of mg/L. |  |  |
| - Select specific units on OCMA-350 for oil in water analysis |  |  |
| - Accurately pipette sample and Solvent S-316 in 40ml vial |  |  |
| - Select the measurement range for the sample under test and take the appropriate volume of sample as per manual instruction. |  |  |
| - Perform sample extraction with vigorous shaking for break between the solvent on bottom and water on the top. |  |  |
| - In case of emulsion treat the extract to break it. Carefully pipette the as much solvent/extract as possible without taking the water layer in a beaker. |  |  |
| - Break the emulsion by adding specific chemical Sodium Sulfate by stirring with glass rod |  |  |
| - Put a conditioned 11cm No. 40 whatman filter paper into a glass funnel and filter the extract to remove the Na2SO4 |  |  |
| - If less than 6ml of solvent/extract are recovered, add a measured amount of solvent, and account for dilution in the results. |  |  |
| - After pipetting 6ml of solvent/extract fill and place the OCMA-350 cell into equipment for measurement. |  |  |
| - For measurement of Total Petroleum Hydrocarbon (TPH). Treat the extract with conditioned Silica Gel, filtered with No. 40 Whatman filter paper and premeasure. |  |  |
| - The difference, Oil and Grease minus Total Petroleum Hydrocarbon, is the non-petroleum Hydrocarbon. |  |  |
| - Perform the zero and span calibration before the measurement. |  |  |
| - Record and report the Results with proper units and uncertainty. |  |  |
| - Close the Cover on the unit to avoid dust. |  |  |
| - When filling the cell, make sure there are no bubbles in the cell. |  |  |
| - Do not touch the ends of a cell. |  |  |
| - Do not spill the solvent on the measuring unit during cell insertion. |  |  |
| - For accurate measuring results, calibration and measurement should be started after the “warmup” lamp off. |  |  |
| - Wear gloves when handling it and avoid inhaling of vapors or ingest liquid. |  |  |

Candidate’s Signature__________________ Assessor’s Signature____________________

Date: _______________________________

Assessors Judgment Guide

| Qualification | National Diploma- Level-5 in “Water Quality and Resource Management” (Technician-Supervisor) |
| --- | --- |
| Competency Standard | Perform Testing/Determination of Oil and Grease in Wastewater |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: ______________________Signature: _______________ |
| Assessment Outcome | COMPETENT NOT YETCOMPETENT Name of the Assessor________________________________________________ Assessor’s code: ____________________________________________________ Signature: ________________________________ |

| Assessment Summary (to be filled by the assessor) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Activity | Method |  |  |  |  | Result |  |
| Nature of Activity | Written | Oral | Observation | Portfolio | Role Play | Competent | Not Yet Competent |
| Practical Skill Demonstration |  |  |  |  |  |  |  |
| Knowledge Assessment |  |  |  |  |  |  |  |
| Other Requirement |  |  |  |  |  |  |  |

Observation Checklist

| Assessment Task |  | - Identify the appropriate test method, Prepare samples and stock solution for testing - Prepare Equipment/Apparatus for testing - Perform tests on samples as per SOPs - Perform Quality Control Checks - Record the results/ Finalize the work - Adopt the precautions during testing work |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| During the practical assessment, candidate demonstrated the following: |  |  |  | Yes | No | Remarks |
|  | Identify appropriate test method, Glassware equipment and safety requirements. |  |  |  |  |  |
|  | Maintain the temperature of sample and working environment before analysis. |  |  |  |  |  |
|  | Maintain the pH of sample. |  |  |  |  |  |
|  | Maintain the pH to <=2 by adding Hydrochloric acid |  |  |  |  |  |
|  | Shake the glass sample bottle to assure good mixing of sample |  |  |  |  |  |
|  | Prepare Zero liquid, span liquid and OCB Standard for Calibration, Pure Solvent S-315 is used as zero solvent. |  |  |  |  |  |
|  | Set up test equipment OCMA-350 and reagents in accordance with the specified work instructions. |  |  |  |  |  |
|  | Perform calibration, zero, span and measurement of zero solvent in terms of mg/L. |  |  |  |  |  |
|  | Select specific units on OCMA-350 for oil in water analysis |  |  |  |  |  |
|  | Accurately pipette sample and Solvent S-316 in 40ml vial |  |  |  |  |  |
|  | Select the measurement range for the sample under test and take the appropriate volume of sample as per manual instruction. |  |  |  |  |  |
|  | Perform sample extraction with vigorous shaking for break between the solvent on bottom and water on the top. |  |  |  |  |  |
|  | In case of emulsion treat the extract to break it. Carefully pipette the as much solvent/extract as possible without taking the water layer in a beaker. |  |  |  |  |  |
|  | Break the emulsion by adding specific chemical Sodium Sulfate by stirring with glass rod |  |  |  |  |  |
|  | Put a conditioned 11cm No. 40 whatman filter paper into a glass funnel and filter the extract to remove the Na2SO4 |  |  |  |  |  |
|  | If less than 6ml of solvent/extract are recovered, add a measured amount of solvent, and account for dilution in the results. |  |  |  |  |  |
|  | After pipetting 6ml of solvent/extract fill and place the OCMA-350 cell into equipment for measurement. |  |  |  |  |  |
|  | For measurement of Total Petroleum Hydrocarbon (TPH). Treat the extract with conditioned Silica Gel, filtered with No. 40 Whatman filter paper and premeasure. |  |  |  |  |  |
|  | The difference, Oil and Grease minus Total Petroleum Hydrocarbon, is the non-petroleum Hydrocarbon. |  |  |  |  |  |
|  | Perform the zero and span calibration before the measurement. |  |  |  |  |  |
|  | Record and report the Results with proper units and uncertainty. |  |  |  |  |  |
|  | Close the Cover on the unit to avoid dust. |  |  |  |  |  |
|  | When filling the cell, make sure there are no bubbles in the cell. |  |  |  |  |  |
|  | Do not touch the ends of a cell. |  |  |  |  |  |
|  | Do not spill the solvent on the measuring unit during cell insertion. |  |  |  |  |  |
|  | For accurate measuring results, calibration and measurement should be started after the “warmup” lamp off. |  |  |  |  |  |
|  | Wear gloves when handling it and avoid inhaling of vapors or ingest liquid. |  |  |  |  |  |
| Competent |  |  | Not Yet Competent |  |  |  |

Knowledge Assessment

| Qualification | National Diploma- Level-5 in “Water Quality and Resource Management” (Technician-Supervisor) |
| --- | --- |
| Competency Standard | Testing/Determination of Oil and Grease in Wastewater |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: _________________ Candidate Signature: ___________ |
| Assessment Outcome | COMPETENT NOT YETCOMPETENT Name of the Assessor: ________________________________________________ Assessor’s code: ____________________________________________________ Signature of the Assessor: _______________________ |

Candidate’s response is not required to be identical, but similar concepts and/or keywords must be used. Oral questioning may be used to clarify candidate understanding of topic and its application.

| Questions (Candidate confidently answered questions correctly and demonstrated understanding of the topics and their application) |  | Satisfactory | Not Satisfactory |
| --- | --- | --- | --- |
|  | Write the working principle of oil and grease measurement in water by OCMA-350 |  |  |
|  | Describe the Infrared Spectroscopic absorption measurement principle |  |  |
|  | What is the role of S-316 (Chlorofluoroethylene) solvent in oil and grease measurement |  |  |
|  | Why filter paper is recommended to desiccate after drying. |  |  |
|  | What are polar and non-polar solvents? Give any 2 examples? |  |  |
|  | Define the procedure of preparation of OCB Standard |  |  |

Feedback to the Candidate

Candidate’s Signature__________________ Assessor’s Signature _________________