National Vocational Certificate Level-4 in

“Water Quality and Resource Management” (Lab Analyst)

(Curriculum)

National Vocational and Technical Training Commission (NAVTTC)

Government of Pakistan

Table of Contents

Introduction	4

a.	Definition/Description of training program	4

b.	Purpose of the training program:	4

c.	Overall objectives of training program:	5

d.	Competencies to be gained after completion of course:	5

e.	Possible available job opportunities, available immediately and later in the future:	6

f.	Trainee entry level:	7

g.	Minimum qualification of trainer:	7

h.	Recommended trainer: trainee ratio	8

i.	Medium of instruction i.e. language of instruction:	8

j.	Duration of the course (Total time, Theory & Practical time):	8

Summary template-overview of the curriculum:	8

Sequence of modules:	13

Module 1021CS-057: Perform Testing of Carbonate (CO3) by Titrimetric Method	20

Module 1021CS-059: Perform Testing of Chlorides (Cl) by Titrimetric/Argentometric  Method	25

Module 1021CS-060. Perform Testing/Determination of Fluoride (F) by ISE Meter	33

Module 1021CS-061. Perform Testing/Determination of Chlorine (Cl2) by DPD/Chlorine Meter	39

Module 1021CS-62: Perform Testing for Total Plate Count (Method-III)	45

Module 1021CS-63: Perform Testing for Total Coliforms and E-coli (IDEXX Method, Method-IV)	52

Module 1021CS-064.	Analyze Method Blank	60

Module 1021CS-065.	Perform Calibration and Calibration Verification	63

Module 1021CS-066.	Analyze Laboratory Control Samples (LCS)	67

Module 1021CS-067.	Analyze Audit Blind Sample	71

Module 1021CS-068.	Perform Correctness of Results	74

Module 1021CS-069.	Determine the Method Detection Limits	79

Module 1021CS-070.	Develop Control Charts	83

Module: 1021 CS-071 Identify Water Resources of Pakistan.	89

Module: Manage meetings with customers	93

Module: Manage workforce planning	96

Module: Undertake project work	101

Module: Identify and communicate trends in career development	106

Module: Apply interpersonal skills	106

Module: Maintain professionalism in the workplace	117

List of Tool & Equipment:	121

# Introduction

# Definition/Description of training program

Water quality and treatment is process that improves the quality of water to make it more acceptable for a specific end-use. The end use may be drinking, industrial water supply, irrigation, river flow maintenance, water recreation or many other uses, including being safely returned to the environment. Water treatment removes contaminants and undesirable components, or reduces their concentration so that the water becomes fit for its desired end-use. This treatment is crucial to human health and allows humans to benefit from both drinking and irrigation use.

In fact, no sphere of life may be identified that does not include the contribution of water technology. Thus, the importance of this diploma may be determined according to its usefulness in our daily life. Therefore, industry expectations for skilled workforce are also dynamic which can only be managed through setting relevant competency standards in collaboration with the leading industries. Being cognizant of this fact, National Vocational & Technical Training Commission (NAVTTC) developed competency standards for water quality and treatment Technology under National Vocational Qualifications Framework (NVQF). These competency standards have been developed by a Qualifications Development Committee (QDC) and validated by the Qualifications Validation Committee (QVC) having representation from the leading department of water industry (WASA, PCSIR, PHED) of the country.

# Purpose of the training program:

The purpose of these qualification is to set high professional standards for Water Quality and Resource Management.

# Overall objectives of training program:

The Overall objectives of the training program are as under:

1.	Improve the professional competence of the trainees

2.	Provide opportunities for recognition of skills attained through non-formal or informal pathways

3.	 Improve the quality and effectiveness of training and assessment for automobile industry

4.	 Enable the existing workforce to capacitate themselves in new water technologies and method

# Competencies to be gained after completion of course:

At the end of the course, the trainee has attained the following core competencies:

- Perform Testing of Bicarbonate (HCO3) by Titrimetric Method

- Perform Testing of Carbonate (CO3) by Titrimetric Method

- Perform Testing of Chlorides (Cl) by Titrimetric/Argentometric Method

- Perform Testing/Determination of Fluoride (F) by ISE Meter

- Perform Testing/Determination of Chlorine (Cl2) by DPD/Chlorine Meter

- Perform Testing for Total Plate Count (Method-III)

- Perform Testing for Total Coliforms and E-coli (IDEXX Method, Method-IV)

- Analyze Method Blank

- Perform Calibration and Calibration Verification

- Analyze Laboratory Control Samples (LCS)

- Analyze Audit Blind Sample

- Perform Correctness of Results

- Determine the Method Detection Limits

- Develop Control Charts

- Identify Water Resources of Pakistan.

# Possible available job opportunities, available immediately and later in the future:

Possible Careers

- Water Hygiene Assistant

- Hydrology Assistant

- Water Quality Surveyor

- Quality Control Technician

- Water Resources Technician

- Water Supply Technician

- Waste Water Testing Technician

- Water Treatment Technologist

- Hydraulics and Irrigation Associate

- Water Works Supervisor

# Trainee entry level:

The entry levels of National Vocational Certificate for “Water Quality and Resource Management” are given below:

| Title | Entry requirements |
| --- | --- |
| National Vocational Certificate Level-4 for “Water Quality and Resource Management” | National Vocational Certificate Level-3 for “Water Quality and Resource Management” |

# Minimum qualification of trainer:

    - Must be a holder of Level 5 Diploma in Water Quality and Resource Management with at least 2 years relevant experience

OR

    - BSc (4 years) in Chemistry or equivalent

OR

    - BSc Engineering Technology (Civil) / B.E Civil /BSc Civil Engineering

# Recommended trainer: trainee ratio

The recommended maximum trainer: trainee ratio for this program is 1 trainer for 25 trainees.

# Medium of instruction i.e. language of instruction:

Instructions will be in Urdu/ English/ Local language.

# Duration of the course (Total time, Theory & Practical time):

Proposed Course Duration - 1200 hrs

## Summary template-overview of the curriculum:

Following is the structure of the course:

| NVQF (Competency Standards) | Level | Category | Estimated Hours |  |  | CH |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Th. | Pr. | T H |  |
| Level 4 |  |  |  |  |  |  |  |
| 11. Water and waste water Testing (Physico-chemical Parameter)-II Water Testing Analyst (Chemical) |  |  |  |  |  |  |  |
| Perform Testing/Determination of Bicarbonate (HCO3) by Titrimetric Method | 4 | Technical | 15 | 62 | 77 | 7.7 |  |
| Perform Testing/Determination of Carbonate (CO3) by Titrimetric Method | 4 | Technical | 15 | 62 | 77 | 7.7 |  |
| Perform Testing/Determination of Chloride (Cl) by Titrimetric/Argentometric Method | 4 | Technical | 15 | 62 | 77 | 7.7 |  |
| Perform Testing/Determination of Chlorine (Cl2) by DPD/Chlorine Meter Method | 4 | Technical | 15 | 62 | 77 | 7.7 |  |
| Perform Testing/Determination of Fluoride (F) by ISE Meter | 4 | Technical | 15 | 63 | 78 | 7.8 |  |
| Total |  |  | 75 | 314 | 386 | 38.6 |  |
| 12. Water Quality Testing for Microbiological Contamination-II Water Testing Analyst (Microbiological) |  |  |  |  |  |  |  |
| Perform Testing for Total Plate Count (Method-III) | 4 | Technical | 20 | 60 | 80 | 8 |  |
| Perform Testing for Total Coliforms and E-coli (Method-IV) | 4 | Technical | 20 | 60 | 80 | 8 |  |
| Total |  |  | 40 | 120 | 160 | 16 |  |
| 13. Quality Control Operations (Quality Control Technician) |  |  |  |  |  |  |  |
| Analyze Method Blank | 4 | Technical | 15 | 60 | 75 | 7.5 |  |
| Perform Calibration and Calibration Verification | 4 | Technical | 15 | 60 | 75 | 7.5 |  |
| Analyze Laboratory Control Samples (LCS) | 4 | Technical | 15 | 60 | 75 | 7.5 |  |
| Analyze Audit Blind Sample | 4 | Technical | 15 | 60 | 75 | 7.5 |  |
| Perform Correctness of Results | 4 | Technical | 15 | 65 | 80 | 8 |  |
| Determine the Method Detection Limits | 4 | Technical | 15 | 60 | 75 | 7.5 |  |
| Develop Control Charts | 4 | Technical | 15 | 50 | 65 | 6.5 |  |
| Total |  |  | 105 | 415 | 520 | 52 |  |
| 14. Water Resources of Pakistan (Water Resources Technician) |  |  |  |  |  |  |  |
| Identify Water Resources of Pakistan | 4 | Technical | 6 | 20 | 26 | 2.6 |  |
| Total |  |  | 6 | 20 | 26 | 2.6 |  |
| 15. Soft skills |  |  |  |  |  |  |  |
| Manage the meetings with customers | 4 | Generic | 6 | 12 | 18 | 1.8 |  |
| Manage workforce planning | 4 | Generic | 6 | 12 | 18 | 1.8 |  |
| Undertake project work | 4 | Generic | 6 | 12 | 18 | 1.8 |  |
| Apply interpersonal skills | 4 | Generic | 6 | 12 | 18 | 1.8 |  |
| Work safely in an office environment | 4 | Generic | 6 | 12 | 18 | 1.8 |  |
| Maintain professionalism in the workplace | 4 | Generic | 6 | 12 | 18 | 1.8 |  |
| Total |  |  | 36 | 72 | 108 | 10.8 |  |
| Grand Total |  |  | 262 | 938 | 1200 | 120.0 |  |

## Sequence of modules:

Module 1021CS-057: Perform Testing of Bicarbonate (HCO3) by Titrimetric Method

Objective: After the completion of this Module, the trainee will be able to develop skills and knowledge required for determination of Bicarbonate (HCO3) in water by Titrimetric  method.

Duration:  77 Hours			Theory: 15 Hours		Practice:  62 Hours                  Credit Hours:  7.7

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Prepare samples for testing | Trainee will be able to: - Check the sample and environmental conditions required for analysis of Alkalinity - Recognize appropriate test method and Glassware and safety requirement - Check for the availability or Prepare the Hydrochloric acid (HCl) Solution as per requirement | - Define the Alkalinity due to Bicarbonate (HCO3) - Enlist the glassware needed for Titration Practical Activity: Prepare 0.02N HCl Solution from 0.1N HCl | Theory- 02 Hrs. Practical- 10 Hrs. Total- 12 Hrs. | - Burette - HCl - Deionized water - Glass ware (Class A) - Auto-pipette | Class Room/lab |
| LU2. Prepare for laboratory/field testing | Trainee will be able to: - Preparation of washing acid - Proper washing of all glassware with washing acid and deionized water - Set up the glassware and filling of burette with HCl - Conduct standardization of 0.02N HCl against 0.05N Na2CO3 | - Explain the procedure of glassware washing - Define the Standardization procedure for HCl. Practical Activity: Standardize the 0.02N HCl with 0.05N Sodium Carbonate (Na2CO3) | Theory- 02 Hrs. Practical- 10 Hrs. Total- 12Hrs. | - Washing acid - Burette - Glass Beaker (Class A) - HCl (0.02 N) - 0.05N Na2CO3 - Deionized water - Auto-pipette - Burette | Class Room Lab |
| LU3. Perform tests on samples | Trainee will be able to: - Prepare the sample by taking 10 sample, few drops of mixed indicator with the appearance of green color - Select the burette for low and high alkalinity range. 25 ml burette with 0.05ml resolution for titration, for low alkalinity less than 50mg/l 10ml burette with 0.02ml resolution - Titrate the prepared sample against 0.02N HCl with the change of color from green to light pink end point - The Bicarbonate is calculated by using following formula HCO3 Alkalinity, mg CaCO3/L; =; A × N × 50000. ml sample. Where: A = ml standard acid used N = Normality of the standard acid. - Analyze Three replicate for each sample. - Store unused reagents and dispose of wastes as required by relevant regulations and codes. Clean and store equipment | - Describe the basic principle of HCO3 Alkalinity Titration - Explain the steps involved in Titration procedure - State the formula for HCO3 Alkalinity measurement Practical Activity: Analyse the sample for Bicarbonate (HCO3) Alkalinity | Theory- 03 Hrs. Practical- 14 Hrs. Total- 17 Hrs. | - Burette - Hydrochloric Acid solution (0.1N) - Hydrochloric Acid solution (0.02N) - Sodium carbonate solution (0.05N) - Bromocresol green + Methyl red (Mixed indicator) - Deionized water - Glass Ware (Class A) - Auto-pipette - Wash Bottle Auto-pipette | Class Room Lab |
| LU4. Perform Quality Control Checks | Trainee will be able to: - Method blank analysis after every 10 samples - Laboratory Control samples of 5 and 100 ppm must be analyzed with every batch of 10 samples. - Check repeatability by reanalyzing a pre-analyzed sample once a day - Record of Daily quality control data in QC Data sheet - Generation of control charts of the recorded data | - Define the Lab Control (LCS) & Audit blind sample (ABS) - Define Quality control practices - Explain repeatability Practical Activity: Analyse the Audit blind sample for HCO3 Alkalinity | Theory- 03 Hrs. Practical- 10 Hrs. 13 Total- 17 Hrs. | - Burette - Hydrochloric Acid solution (0.1N) - Hydrochloric Acid solution (0.02N) - Sodium carbonate solution (0.05N) - Bromocresol green + Methyl red (Mixed indicator) - Deionized water - Glass Ware (Class A) - Auto-pipette - Wash Bottle Auto-pipette | Class Room Lab |
| LU5. Record the results/finalise work | Trainee will be able to: - Note down the Results on analyst workbook. - Calculate the uncertainty of the results - Record and reporting of the results. - Clear and restore work area. | - Enlist the steps for uncertainty measurement. - Define replicate analysis Practical Activity: Record the results on workbook and finalize work | Theory- 03 Hrs. Practical- 10 Hrs. Total- 13 Hrs. . | - Note book - Result record form - Calculator | Class Room Lab |
| LU6. Adopt the precautions during testing work | Trainee will be able to: - Drying of Sodium Carbonate before making solution at 250oC for 4 hours. - Titration must be done carefully and quickly to avoid loss of dissolved gasses during titration. - Dissolved gases can, directly or indirectly contribute to alkalinity; therefore the gases must neither be lost nor gained during sampling, storage and titration. | - Describe the condition required for Sodium Carbonate solution preparation - Define the contribution of dissolved gases in Alkalinity Practical Activity: Perform practical Step to avoid error during the analysis | Theory- 02 Hr. Practical- 08 Hrs. Total- 10 Hrs | - Oven - Sodium Carbonate | Class Room Lab |

## Module 1021CS-057: Perform Testing of Carbonate (CO3) by Titrimetric Method

Objective: After the completion of this Module, the trainee will be able to develop skills and knowledge required for determination of Carbonate (CO3) in water by Titrimetric  method.

Duration:  77 Hours			Theory: 15 Hours		Practice:  62 Hours                  Credit Hours:  7.7

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Prepare samples for testing | Trainee will be able to: - Check the sample and environmental conditions required for analysis of Alkalinity - Recognize appropriate test method and Glassware and safety requirement - Check for the availability or Prepare the Hydrochloric acid (HCl) Solution as per requirement | - Define the Alkalinity due to Carbonate (CO3) - Enlist the glassware needed for Titration Practical Activity: Prepare 0.02N HCl Solution from 0.1N HCl | Theory- 02 Hr. Practical- 10 Hrs. Total- 12 Hrs | - Burette - HCl - Deionized water - Glass ware (Class A) - Auto-pipette | Class Room/lab |
| LU2. Prepare for laboratory/field testing | Trainee will be able to: - Preparation of washing acid - Proper washing of all glassware with washing acid and deionized water - Set up the glassware and filling of burette with HCl - Conduct standardization of 0.02N HCl against 0.05N Na2CO3 | - Describe the preparation of washing acid - Define the procedure for preparation of 0.02N HCl from 0.1N HCl Practical Activity: Standardize the 0.02N HCl with 0.05N Sodium Carbonate (Na2CO3) | Theory- 02 Hr. Practical- 9 Hrs. Total- 11 Hrs | - Washing acid - Burette - Glass Beaker (Class A) - HCl (0.02 N) - 0.05N Na2CO3 - Deionized water - Auto-pipette - Burette | Class Room Lab |
| LU3. Perform tests on samples | Trainee will be able to: - Prepare sample by taking 10 ml sample add 2 or 3 drops of phenolphthalein pink color indication shows the pH range 8.3 or above. - Titrate against HCl with 10ml burette of 0.02ml resolution as per standard. - Color changes from pink to transparent color. - Alkalinity determined is called as Phenolphthalein Alkalinity (P) as it gives color at pH 8.3. Carbonate Alkalinity as CaCO3 mg/L = 2P - Analyze three replicate for each sample. - Store unused reagents and dispose of wastes as required by relevant regulations and codes. - Clean and store equipment | - Describe the basic principle of CO3 Alkalinity Titration - Explain the steps involved in Titration procedure - Explain Phenolphthalein Alkalinity (P) Practical Activity: Analyse the sample for Carbonate (CO3) Alkalinity | Theory- 03 Hrs. Practical- 17 Hrs. Total- 20 Hrs | - Burette - Hydrochloric Acid solution (0.1N) - Hydrochloric Acid solution (0.02N) - Sodium carbonate solution (0.05N) - Bromocresol green + Methyl red (Mixed indicator) - Deionized water - Glass Ware (Class A) - Auto-pipette - Wash Bottle Auto-pipette | Class Room Lab |
| LU4. Quality Control Checks | Trainee will be able to: - Method blank analysis after every 10 samples - Laboratory Control samples of must be analyzed with every batch of 10 samples. - Check repeatability by reanalyzing a pre-analyzed sample once a day - Record of Daily quality control data in QC Data sheet - Generation of control charts of the recorded data | - Define the Lab Control (LCS) & Audit blind sample (ABS) - Define Quality control practices - Explain repeatability Practical Activity: Analyse the Audit blind sample for CO3 Alkalinity | Theory- 03 Hrs. Practical- 14 Hrs. Total- 17 Hrs. | - Burette - Hydrochloric Acid solution (0.02N) - Lab Control Sample - Deionized water - Glass Ware (Class A) - Auto-pipette - Wash Bottle | Class Room Lab |
| LU5. Record the results/finalise work | Trainee will be able to: - Note down the Results on analyst workbook. - Calculate the uncertainty of the results - Record and reporting of the results. - Clear and restore work area. | - Enlist the steps for uncertainty measurement. - Define replicate analysis Practical Activity: Record the results on workbook and finalize work | Theory- 03 Hrs. Practical- 07 Hrs. Total- 10 Hrs.. | - Note book - Result record form - Calculator | Class Room Lab |
| LU6. Adopt the precautions during testing work | Trainee will be able to: - Drying of Sodium Carbonate before making solution at 250oC for 4 hours. - Titration must be done carefully and quickly to avoid loss of dissolved gasses during titration. - Dissolved gases can, directly or indirectly contribute to alkalinity; therefore the gases must neither be lost nor gained during sampling, storage and titration. | - Describe the condition required for Sodium Carbonate solution preparation - Define the contribution of dissolved gases in Alkalinity Practical Activity: Perform practical Step to avoid error during the analysis | Theory- 02 Hr. Practical- 05 Hrs. Total- 07 Hrs | - Oven - Sodium Carbonate | Class Room Lab |

## Module 1021CS-059: Perform Testing of Chlorides (Cl) by Titrimetric/Argentometric  Method

Objective: After the completion of this Module, the trainee will be able to develop skills and knowledge required for determination of Chlorides (Cl) in water by Argentometric Titration method

Duration:  77 Hours			Theory: 15 Hours		Practice:  62 Hours                  Credit Hours:  7.7

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Prepare samples for testing | Trainee will be able to: - Check the sample and environmental conditions required for analysis of Alkalinity - Recognize appropriate test method and Glassware and safety requirement - Check for the availability or Prepare the 0.0141N Silver Nitrate (AgNO3) Solution as per requirement | - Define the Argentometric Titration - Define the permissible limit for Chloride - Enlist the glassware needed for Titration Practical Activity: Prepare 0.0141N AgNO3 Solution | Theory- 02 Hr. Practical- 10 Hrs. Total- 12 Hrs. | - Burette - AgNO3 - Deionized water - Glass ware (Class A) - Auto-pipette | Class Room/lab |
| LU2. Prepare for laboratory/field testing | Trainee will be able to: - Preparation of washing acid - Proper washing of all glassware with washing acid and deionized water - Set up the glassware and filling of burette with AgNO3 - Conduct standardization of 0.0141N AgNO3 Solution against Sodium Chloride (NaCl) | - Explain the procedure of glassware washing - Define the Standardization procedure for 0.0141N AgNO3 Solution Practical Activity: Standardize the 0.0141N AgNO3 Solution with 0.0141N Sodium Chloride(NaCl) and calculate multiplication factor | Theory- 02 Hrs. Practical- 09 Hrs. Total- 11 Hrs. | - Washing acid - Burette - Glass Ware (Class A) - 0.0141N AgNO3 Solution - 0.0141N NaCl - Potassium Chromate (K2CrO4) - Deionized water - Auto-pipette | Class Room Lab |
| LU3. Perform tests on samples | Trainee will be able to: - Prepare sample by Taking 20 ml Sample in beaker and few drops of K2CrO4 indicator in it. - Titrate the prepared sample against 0.0141N AgNO3 with 25ml burette of 0.05ml resolution The burette reading is then noted when sample color changes to pinkish yellow - Calculation: mg Cl -1/L=; (A – B) × N ×1000 ×35.6. Where: A = mL titration for sample B = mL titration for blank C = Normality of AgNO3; ml of sample. - Analyze Three replicate for each sample. - Store unused reagents and dispose of wastes as required by relevant regulations and codes. - Clean and store equipment | - Describe the basic principle of Chloride testing - Enlist the steps involved in Titration procedure - State the formula for Chloride measurement Practical Activity: Analyse the sample for Chloride | Theory- 03 Hrs. Practical- 17 Hrs. Total- 20 Hrs. | - Burette - Glass Ware (Class A) - 0.0141N AgNO3 Solution - 0.0141N NaCl - Potassium Chromate - (K2CrO4) - Deionized water - Auto-pipette | Class Room Lab |
| LU4. Quality Control Checks | Trainee will be able to: - Method blank analysis after every 10 samples - Laboratory Control samples of 5 and 100 ppm must be analyzed with every batch of 10 samples. - Check repeatability by reanalyzing a pre-analyzed sample once a day - Record of Daily quality control data in QC Data sheet - Generation of control charts of the recorded data | - Define Method blank - Define Quality control practices Practical Activity: Analyse the Audit blind sample for Chloride | Theory- 03 Hrs. Practical- 14 Hrs. Total- 17 Hrs. | - Burette - Glass Ware (Class A) - 0.0141N AgNO3 Solution - 0.0141N NaCl - Potassium Chromate (K2CrO4) - Deionized water - Auto-pipette - Lab Control Sample | Class Room Lab |
| LU5. Record the results/finalise work | Trainee will be able to: - Note down the Results on analyst workbook. - Calculate the uncertainty of the results - Record and reporting of the results. - Clear and restore work area. | - Enlist the steps for uncertainty measurement. - Describe the Trend analysis of Control Chart Practical Activity: Generate Control chart of given data | Theory- 03 Hrs. Practical- 07 Hrs. Total- 10 Hrs. | - Note book - Result record form - Calculator | Class Room Lab |
| LU6. Adopt the precautions during testing work | Trainee will be able to: - Store Silver nitrate stock solution in amber colored bottle. - The glass ware should be properly washed and cleaned before analysis - Careful weighing of silver nitrate. - Ensure the sample should be free of interfering radicals otherwise their presence may introduce errors in the chloride estimation or may dilute the sample to avoid interference | - Describe the storage condition for AgNO3 solution Practical Activity: Perform analysis on high chloride concentration sample and report result | Theory- 02 Hr. Practical- 05 Hrs. Total- 07 Hrs | - Sample - Glassware - Burette - 0.0141N AgNO3 - Potassium Chromate (K2CrO4) | Class Room Lab |

## Module 1021CS-060. Perform Testing/Determination of Fluoride (F) by ISE Meter

Objective: After the completion of this module, the Trainee will be able to develop expertise and competence required to prepare standards/samples for testing of Fluoride, Quality Control (QC) checks, data recording/reporting of results  and necessary precautions required during testing.

Duration:  78 Hours			Theory: 15 Hours		Practice: 63 Hours                  Credit Hours: 7.8 Hours

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Prepare samples for testing | Trainee will be able to: - Identify the samples and environmental condition required for testing - Select appropriate test method, glassware and equipment - Prepare the required standards for verification of test methods/ calibration of equipments. - Adopt the appropriate safety precautions | - Describe the environmental conditions and safety precautions required during testing - Explain the preparation of standards as per reference procedure. - Enlighten the verification of test method Practical Activity: Prepare the Flouride standard of given concentration | Theory- 2 Hrs Practical- 11 Hrs Total- 12 Hrs | - ISE Meter - Auto Pipette - Fluoride Standard - Fluoride Buffer - Glass ware | Class Room Lab |
| LU2. Prepare for laboratory testing | Trainee will be able to: - Set up test equipment and reagents in accordance with the specified work instructions - Conduct pre-use and safety checks - Check the calibration status of equipment. Perform calibration if required | - Enlist the regents require during testing - Explain the QC checks involve in determination of Flouride - Describe the calibration procedure for ISE meter Practical Activity: Calibrate the meter with given standards | Theory- 2 Hrs Practical-9 Hrs Total- 11 Hrs | - ISE Meter - Auto Pipette - Fluoride Standard - Fluoride Buffer - Glass ware | Class Room Lab |
| LU3. Perform tests on samples | Trainee will be able to: - Take volume as per standard of deionized water in a cuvet and add specific volume of fluoride Buffer (pH: 5.3-5.5) in it. - Dip the electrode in the solution with constant stirring. This will show zero concentration - Note the concentration of known standards for verification of calibration and efficiency of electrode. - Analyze three replicate for each sample. - Store unused reagents and dispose of wastes as required by relevant regulations and codes. - Clean and store equipment | - Explain the working principle of ISE meter - State the range of Fluoride buffer - Define the zero concentration - Describe the verification of calibration Practical Activity: Demonstrate the working of ISE meter by Analyzing the given sample | Theory- 3 Hrs Practical- 17 Hrs Total- 20 Hrs | - ISE Meter - Auto Pipette - Fluoride Standard - Fluoride Buffer - Glass ware | Class Room Lab |
| LU4. Quality Control Checks | - Analyze method blank after every 10 samples - Analyze Laboratory control sample with every batch of 10 samples. - Check repeatability by reanalyzing a pre-analyzed sample once a day - Recorded Quality control data in QC Data sheet. - Generate the control chart of laboratory control sample | - Explain the repeatability and its procedure - State the method blank - Elaborate the trend analysis of given control charts Practical Activity: Generate the control chart from provided data | Theory- 3 Hrs Practical- 14 Hrs Total- 17 Hrs | - ISE Meter - Auto Pipette - Fluoride Standard - Fluoride Buffer - Glass ware | Class Room Lab |
| LU5. Record the results/ Finalize work | - Note down the results on analyst workbook - Calculate uncertainty of the results - Record the results on result record form and submit to reporting section - Clear and restore work area | - Enlist the steps involve in measurement uncertainty - Describe the replicate analysis Practical Activity: Calculate the uncertainty of given data | Theory- 3 Hrs Practical- 7 Hrs Total- 10 Hrs | - ISE Meter - Auto Pipette - Fluoride Standard - Fluoride Buffer - Glass ware | Class Room Lab |
| LU6. Adopt the precautions during testing work | - Check the consistency of the instrument by analyzing known standards. - The pH of the buffer solution must be in range of 5.3-5.5 for the accurate work. - Alkalinity, iron, aluminum and free chlorine may interfere during the analysis of fluoride. - It should be kept in mind that at alkaline pH the hydroxide ions also interfere the electrode response. - Use samples and standards at the same temperature. | - Enlist the ions that interfere during analysis of Fluoride - Describe the pH at which hydroxide ions interfere - Describe the permissible limit of Fluoride in drinking water Practical Activity: Analyze the given sample for Fluoride test | Theory- 2 Hrs Practical- 5 Hrs Total- 7 Hrs | - ISE Meter - Auto Pipette - Fluoride Standard - Fluoride Buffer - Glass ware | Class Room Lab |

## Module 1021CS-061. Perform Testing/Determination of Chlorine (Cl2) by DPD/Chlorine Meter

Objective: After the completion of this module, the Trainee will be able to develop expertise and competence required to prepare standards/samples for testing of Chlorine, Quality Control (QC) checks, data recording/reporting of results  and necessary precautions required during testing.

Duration:  78 Hours			Theory: 15 Hours		Practice: 63 Hours                  Credit Hours: 7.8 Hours

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Prepare samples for testing | Trainee will be able to: - Identify the samples and environmental condition required for testing - Select appropriate test method, glassware and equipment - Prepare the required standards for verification of test methods/ calibration of equipments. - Adopt the appropriate safety precautions | - Describe the environmental conditions and safety precautions required during testing - Explain the preparation of standards as per reference procedure. - Describe the verification of test method Practical Activity: Prepare the Chlorine standard of given concentration | Theory- 2 Hrs Practical- 10 Hrs Total- 12 Hrs | - Reagent B for free Chlorine determination - Reagent C for combined chlorine determination - DPD Reagent - Deionized water - Glass ware - Lint free tissue | Class Room Lab |
| LU2. Prepare for laboratory testing | Trainee will be able to: - Set up test equipment and reagents in accordance with the specified work instructions - Conduct pre-use and safety checks - Check the calibration status of equipment. Perform calibration if required | - Enlist the regents require during testing - Describe the calibration procedure for Chlorine meter Practical Activity: Calibrate the meter with given standards | Theory- 2 Hrs Practical-9 Hrs Total- 11Hrs | - Reagent B for free Chlorine determination - Reagent C for combined chlorine determination - DPD Reagent - Deionized water - Glass ware - Lint free tissue | Class Room Lab |
| LU3. Perform tests on samples | Trainee will be able to: - Turn on the meter as per the work instructions. - Select free or total chlorine mode as per requirement of test - Add 5ml of reagent B and 5ml of sample in cuvet and shake gently. - Place the cuvet in to the holder and ensure that the notch on the cap is positioned securely in to the groove. - Press “ZERO” Wait for a few seconds, instrument will be zeroed - Remove the cuvet only in case of total chlorine measurement add three drops of reagent C to the cuvet - Add the contents of one packet of DPD reagent to the cuvet and shake gently. - Place the cuvet in to the holder wait for one minute in case of free chlorine or 2 minutes and thirty seconds in case of total chlorine or as per requirement of method. Run the sample the display will show reading in ppm. - Analyzed three replicate for each sample. - Store unused reagents and dispose of wastes as required by relevant regulations and codes. - Clean and store equipment. | - Explain the working principle of Chlorine meter - Describe the procedure for instrument to zero - Describe the verification of calibration Practical Activity: Demonstrate the working of ISE meter by Analyzing the given sample | Theory- 3 Hrs Practical- 17 Hrs Total- 20 Hrs | - Reagent B for free Chlorine determination - Reagent C for combined chlorine determination - DPD Reagent - Deionized water - Glass ware - Lint free tissue | Class Room Lab |
| LU4. Quality Control Checks | - Analyze method blank after every 10 samples - Analyze Laboratory control sample with every batch of 10 samples. - Check repeatability by reanalyzing a pre-analyzed sample once a day - Recorded Quality control data in QC Data sheet. - Generate the control chart of laboratory control sample | - Explain the repeatability and its procedure - State the method blank - Elaborate the trend analysis of given control charts Practical Activity: Generate the control chart from provided data | Theory- 3 Hrs Practical- 14 Hrs Total- 17 Hrs | - Reagent B for free Chlorine determination - Reagent C for combined chlorine determination - DPD Reagent - Deionized water - Glass ware - Lint free tissue | Class Room Lab |
| LU5. Record the results/ Finalize work | - Note down the results on analyst workbook - Calculate uncertainty of the results - Record the results on result record form and submit to reporting section - Clear and restore work area | - Enlist the steps involved in measurement uncertainty - Describe the replicate analysis Practical Activity: Calculate the uncertainty of given data | Theory- 3 Hrs Practical- 8 Hrs Total- 11 Hrs | - Work book - Result record form - Calculator | Class Room Lab |
| LU6. Adopt the precautions during testing work | - Do not touch the cuvet walls with hands. - Do not let the test sample stand too long after reagent is added accuracy may be lost. - Whenever the cuvet is placed into the measurement cell, it must be free of fingerprints, oil or dirt. - To obtain the representative sample, gently but thoroughly, mix it before sample is taken. - Each time the cuvet is used, tighten the cap to the same degree. | - Define the representative sample - Describe the permissible limit of Chlorine in drinking water Practical Activity: Analyze the given sample for Chlorine test | Theory- 2 Hrs Practical- 5 Hrs Total- 7 Hrs | - Reagent B for free Chlorine determination - Reagent C for combined chlorine determination - DPD Reagent - Deionized water - Glass ware - Lint free tissue | Class Room Lab |

## Module 1021CS-62: Perform Testing for Total Plate Count (Method-III)

Objective: After the completion of this module, the Trainee will be able to develop skill and competence for the preparation of media, testing  and confirmation of total plate count  using Pour Plate Method.

Duration:  80 Hours			Theory:  20 Hours		Practical:  60 Hours                  Credit Hours:  8

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Prepare to undertake basic microbiological tests | Trainee will be able to: - Prepare sterile environment for undertaking tests. - Collect samples for microbiological examination in clean, sterile, non-reactive borosilicate glass or plastic bottle. - Ensure that water sample is appropriate for the test requested as per standards. - When the sample is collected, leave ample air space in the bottle at least 2.5 cm to facilitate mixing by shaking before examination. - Ice samples during transport to the laboratory if they cannot be processed within 1 hour after collection. - Maintain samples in the dark and keep cool with ice at < 8 oC but not frozen. | - Understanding of sterile environment required for microbiology - Knowledge of sample collection and transportation under controlled environment Practical Activity: Demonstrate the sample collection and transportation procedures for microbiological analysis | Theory-5 Hrs Practical-10 Hrs Total-15 Hrs | - surgical mask - gloves - microbiological sample collection kit - coolants in ice cooler - microbiological sample collection sterile bottles - Safety Shoes | Microbiological Laboratory Field |
| LU2. Adopt recommended preparatory method for Total Plate Count | Trainee will be able to: - Use commercially available media for detection of total bacterial count following plate count technique. - Prepare recommended media (Plate count agar) by weighing the amount as per manufacturers’ instruction using calibrated analytical balance. - Sterilize media using autoclave as per standards. - Maintain pH of media after sterilization as per manufacturer’s instruction after sterilization. - Bring medium at room temperature before use or store in the refrigerator. | - Enlist the different steps of media preparation for total plate count method. - Describe the working of analytical balance - Describe the working of pH meter. - Describe the operation of autoclave for sterilization. Practical Activity: Demonstrate media preparation for Total Plate count. | Theory-5 Hrs Practical-20 Hrs Total-25 Hrs | - surgical mask - gloves - Plate Count agar - Analytical balance - Magnetic stirrer - pH meter - Media preparation bottles - De-ionized water - Microwave oven - Autoclave | Microbiological Laboratory |
| LU3. Perform handling or basic testing of samples for Total Plate Count | Trainee will be able to: - Prepare dilutions for polluted water samples. - Pour recommended volume of water sample to be tested into petri plate using auto-pipette. - Pour media into plates, shake well, and allow solidifying and incubating at the recommended temperature for 48 hours. - Count bacterial colonies after incubation and report as CFU/ml as per standards. | - Enlist the different steps for the preparation of dilutions. - Estimate the volume of water sample and pouring in petri plates. - Describe the processing and incubation of water samples using total plate count. - Define bacterial counting and reporting mechanism after incubation. Practical Activity: Demonstrate inoculation of water sample of different dilutions using pour plate method for total plate count. | Theory-5 Hrs Practical-20 Hrs Total-25 Hrs | - surgical mask - gloves - Plate count agar - Test tubes - Media dispenser - Analytical balance - Autoclave - Magnetic stirrer - Media preparation bottles - De-ionized water - Auto-pipette - Auto-pipette tips(1ml) - Petri plates - Mechanical shaker - Biological safety cabinet - Incubator - Colony counter | Microbiological Laboratory |
| LU4. Maintain Environment and Quality Control | Trainee will be able to: - Adopt recommended safety precautions when handling biologically hazardous specimen. - Accept or reject media / sample based on the condition and sterility as per standards. - Check and maintain sterility, temperature and humidity of rooms and equipment as per standards. - Use controls i.e. positive and negative controls along with the tests as per standards. - Clean and store equipment in accordance with laboratory guidelines. - Sterilize equipment as required - Store chemicals, media and controls as per standards. - Clean and disinfect area. - Properly dispose contaminated waste in accordance with the standards. | - Describe the safety guidelines in microbiology laboratory. - Describe media quality control for receiving. - Enlist environmental conditions i.e. temperature and humidity. - Describe the use of controls i.e. +ve/-ve - Understanding of sterilization and safety precautions. - Describe the sterilization technique for equipment and glassware. - Describe the storage protocols for chemical/media and standards. Practical Activity: Demonstrate the performance testing of media for quality control. | Theory-5 Hrs Practical-10Hrs Total-15 Hrs | - surgical mask - gloves - Media:broth/agar - Disposable petriplates - Biological safety cabinet - Bunsen burner - Incubator - Temperature and humidity meter - Thermometer - Autoclave - Ethanol | Microbiological Laboratory |

## Module 1021CS-63: Perform Testing for Total Coliforms and E-coli (IDEXX Method, Method-IV)

Objective: After the completion of this module, the Trainee will be able to develop skill and competence for the handling and  testing  of Total Coliforms and E-coli using IDEXX Method.

Duration:  80 Hours			Theory: 20 Hours		Practical:  60 Hours                  Credit Hours: 8

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Prepare to undertake basic microbiological tests | Trainee will be able to: - Prepare sterile environment for undertaking tests. - Collect samples for microbiological examination in clean, sterile, non reactive borosilicate glass or plastic bottle. - Ensure that water sample is appropriate for the test requested as per standards. - When the sample is collected, leave ample air space in the bottle at least 2.5 cm to facilitate mixing by shaking before examination as per standards. - Ice samples during transport to the laboratory if they cannot be processed within 1 hour after collection. - Maintain samples in the dark and keep cool with ice at < 8 oC but not frozen as per standards. - For coliforms and E-coli analyses, the holding time from collection to analysis is 30 hour. | - Understanding of sterile environment required for microbiology - Knowledge of sample collection and transportation under controlled environment Practical Activity: Demonstrate the sample collection and transportation procedures for microbiological analysis | Theory-5 Hrs Practical-10Hrs Total-15 Hrs | - surgical mask - gloves - microbiological sample collection kit - coolants in ice cooler - microbiological sample collection sterile bottles - Safety Shoes | Microbiological Laboratory Field |
| LU2. Adopt recommended method for handling and testing of samples | Trainee will be able to: - Use commercially available media for detection of coliforms & E-coli following Colilert IDEXX quanti tray method. - Use recommended techniques to prepare samples for testing. - Add recommended volume of water sample in sterile mixing bottle and one package of the colilert reagent. - Transfer sample/reagent mixture from the mixing bottle into a quanti-tray. - Seal tray in IDEXX sealer, label tray on the back with sample code, date and time. - Incubate tray at the temperature set by the standard. - Count the number of small and large positive wells and report results using IDEXX MPN table to find the most probable number for Total Coliform. - Obtain E. coli results by placing the wells under a black light and counting the number of fluorescent wells. | - Define the principle of commercially available media for total coliforms and E-coli. - Enlist the steps of sample preparation before the processing using quanti-tray method. - Describe the handling of water sample using quanti-tray sealer. - Define incubation. - Describe the counting of wells for reporting of results of total coliforms using MPN table. - Describe the procedure of counting of wells for reporting the results of E-coli using MPN table. Practical Activity: Demonstrate the inoculation of water sample for total coliforms and E-coli using IDEXX-colilert quanti-tray method. | Theory-10 Hrs Practical-40 Hrs Total-50 Hrs | - surgical mask - gloves - Plate Count agar - Analytical balance - Magnetic stirrer - pH meter - Media preparation bottles - De-ionized water - Microwave oven - Autoclave | Microbiological Laboratory |
| LU3. Maintain Environment and Quality Control | Trainee will be able to: - Adopt recommended safety precautions when handling biologically hazardous specimen. - Accept or reject media / sample based on the condition and sterility as per standards. - Check and maintain sterility, temperature and humidity of rooms and equipment as per standards. - Use controls i.e. positive and negative controls along with the tests as per standards. - Clean and store equipment in accordance with laboratory guidelines. - Sterilize equipment as required - Store chemicals, media and controls as per standards. - Clean and disinfect area. - Properly dispose contaminated waste in accordance with the standards. | - Describe the safety guidelines in microbiology laboratory. - Describe media quality control for receiving. - Enlist environmental conditions i.e. temperature and humidity. - Describe the use of controls i.e. +ve/-ve - Understanding of sterilization and safety precautions. - Describe the sterilization technique for equipment and glassware. - Describe the storage protocols for chemical/media and standards. Practical Activity: Demonstrate the performance testing of media for quality control. | Theory-5 Hrs Practical-10 Hrs Total-15 Hrs | - surgical mask - gloves - Media:broth/agar - Disposable petriplates - Biological safety cabinet - Bunsen burner - Incubator - Temperature and humidity meter - Thermometer - Autoclave - Ethanol | Microbiological Laboratory |

## Module 1021CS-064.	Analyze Method Blank

Objective: After the completion of this module, the trainee will be able to develop skill and competence required to prepare,analyze and evaluate the method blank sample

Duration:  15 Hours			Theory: 3 Hours		Practice: 12 Hours                  Credit Hours:  1.5

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Prepare Method Blank Sample | - Prepare method blank from reagent water and all reagents that normally are in contact with the sample during the entire analytical procedure - Take the De-ionized water and add all the reagents which are used in the sample preparation | - Define method Blank - Enlist the reagent and glassware required for method blank preparation Practical Activity: Demonstrate method blank preparation | Theory- 1.5Hrs Practical- 02 Hrs Total- 3.5 Hrs | - Type A Glassware - Autopipett - Deionized Water | Lab |
| LU2. Analyze Method Blank | - Analyse Method blank with every analytical batch in order to determine the contribution of reagents and the preparative analytical steps to error in the measurement. - Set the instrument according to equipment instructions - Take an analyse free sample through the analysis using the same reagents, glassware, and instrumentation - Analyze the method blank sample on the instrument at the start of the sample analysis and also at different intervals | - Describe the basic purpose of method blank analysis - Describe various steps involved in method blank analysis Practical Activity: Demonstrate the method blank preparation and analysis | Theory- 0.5 Hrs Practical-3 Hrs Total- 3.5 Hrs | - Type A Glassware - Autopipett - Deionized Water | Lab |
| LU3. Perform the Evaluation | - Evaluates The method blank results for the presence of contamination. - In case of unacceptable contamination present in the method blank, Identify the sources of contaminations and take corrective actions. - Re-prepared and re-analysed the sample as the results is suspected if analyse in the method blank is greater than Method Detection Limit | - Describe the steps require to be taken for contaminated method blank Practical Activity: Evaluate the method blank results. | Theory- 0.5 Hrs Practical-3 Hrs Total- 3.5 Hrs | • Type A Glassware • Autopipett • Deionized Water | Lab |

## Module 1021CS-065.	Perform Calibration and Calibration Verification

Objective: After the completion of this module, the trainee will be able to perform calibration and verification of the instruments

Duration:  15 Hours			Theory: 3 Hours		Practice: 12 Hours                  Credit Hours:  1.5

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Perform instrument calibration | - Prepare the standard solutions - Perform instrument calibration by the use of reference standards or in-house standards. - performs instruments calibration according to manufacturer instructions | - Define calibration - Describe the purpose and importance of instrument calibration Practical Activity: Prepare standard solutions for calibration | Theory- 1.5Hrs Practical- 02 Hrs Total- 3.5 Hrs | - Type A Glassware - Autopipett - Deionized Water - Certified reference material/standards - Instrument instruction manual | Lab |
| LU2. Perform initial calibration (IC) | - Perform initial calibration with in-house standards or standard solutions made from the certified reference materials at levels appropriate for the analysis. - Prepare stock standard solution - Prepare minimum three desired standards solution at levels appropriate for the analysis or calibration range - Set the instrument according to equipment instructions - Analyse these standards solutions on the instrument to generate the standard curve | - Define stock solution - Enlist the minimum standards requirement for initial calibration - Describe the standard curve Practical Activity: Demonstrate standard curve generation | Theory- 1 Hrs Practical-5 Hrs Total- 6 Hrs | - Type A Glassware - Autopipett - Deionized Water - Certified reference material/standards - Instrument instruction manual | Lab |
| LU3. Perform continued calibration verification | - Perform continued calibration verification (ccv) at specified intervals throughout the analytical sequence. - Verified the calibration again through the analysis of independently prepared standard solutions - Also used the pre analysed sample for calibration verification | - Enlist the standard required for verification of calibration Practical Activity: Demonstrate the continued calibration verification | Theory- 0.5 Hrs Practical-5 Hrs Total- 5.5 Hrs | - Type A Glassware - Autopipett - Deionized Water - Certified reference material/standards - Instrument instruction manual | Lab |

## Module 1021CS-066.	Analyze Laboratory Control Samples (LCS)

Objective: After the completion of this module, the Trainee will be able to  Prepare the laboratory control sample for method validation.

Duration:  15 Hours			Theory: 3 Hours		Practice: 12  Hours                  Credit Hours:  1.5

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Prepare the laboratory control sample | - Prepare the laboratory control sample from solution of Certified Reference Material in de-ionized water as per standards. - Prepare first Stock solution from certified standards - Prepare then sub stock standard of desired concentration. - Use fallowing dilution formula for preparation of stock and sub stock standards C1V1 = C2V2 Where C1= initial concentration V1=Volume required C2=concentration required V2=final volume - Use all glassware of Type-A, - Use calibrated autopipette | - Define Lab control sample - Describe the dilution formula for preparation of stock and sub stock standards Practical Activity: Prepare the standard stock and sub stock standards solutions. | Theory- 1Hrs Practical- 4 Hrs Total- 5 Hrs | - Type A Glassware - Autopipett - Deionized Water - Certified reference material/standards - Instrument instruction manual | Lab |
| LU2. Analyze the Lab Control Sample | - Analyze the Lab Control in each batch after every 10 samples. - Plot the result on control chart - If the result of the Lab Control Sample with the limits of control chart go for analysis - If the result of the Lab Control Sample exceeds the limits of control chart take corrective actions | - Describe the evaluation of the result of the lab control sample Practical Activity: Demonstrate the analysis of lab control sample through control chart | Theory- 1 Hrs Practical-4 Hrs Total- 5Hrs | - Type A Glassware - Autopipett - Deionized Water - Certified reference material/standards - Instrument instruction manual | Lab |
| LU3.Corrective Actions | - Check to see if the LCS was prepared and analyzed according to the method. - Check calibration of the instrument - Recalibrate and reanalyze LCS and samples | - Define corrective action - Enlist corrective actions if the result of the lab control sample exceeds the control limit Practical Activity: Plot the result of lab control sample on control chart | Theory- 1 Hrs Practical-4 Hrs Total- 5 Hrs | - Control charts Limits | Lab |

## Module 1021CS-067.	Analyze Audit Blind Sample

Objective: After the completion of this module, the Trainee will be able to  Prepare the audit blind sample for performance evaluation

Duration:  15 Hours			Theory: 3 Hours		Practice: 12 Hours                  Credit Hours:  1.5

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Prepare Audit Blind Sample | - Prepare the Audit blind sample for performance evaluation of lab analysts. - Prepare lab control sample of the desired parameters as fallowing - Prepare the stock standard solution from the certified reference material. - Prepare the desired concentration of the desired parameters from stock solution | - Describe the basic purpose of the Audit blind sample - Describe the dilution formula for preparation of stock and sub stock standards Practical Activity: Prepare lab control sample for the given concentration | Theory- 1 Hrs Practical- 4 Hrs Total- 5 Hrs | - Type A Glassware - Autopipett - Deionized Water - Certified reference material/standards - Instrument instruction manual | Lab |
| LU2.Analyse the Audit Blind Sample | - Set the instrument according to equipment instructions. - Analyse the laboratory control sample as per standard method. - The laboratory analyst analyse the audit blind sample and submits the results | - Describe importance of audit blind sample Practical Activity: Analyze the given audit blind sample | Theory- 1 Hrs Practical-4 Hrs Total- 5 Hrs | - Type A Glassware - Autopipett - Deionized Water - Certified reference material/standards - Instrument instruction manual | Lab |
| LU3. Evaluation of the Audit Blind Sample results | - Calculate per cent recoveries for each parameter. - The per cent recovery is calculated as: - Recovery (%) =found value x 100 True value - Evaluate the results of LCS “SATISFACTORY” (if found in the acceptance range) and “UNSATISFACTORY”. (if not found in the acceptance rangeRecalibrate and reanalyze LCS and samples | - Describe the percent recovery formula - Define ‘found value’ and ‘true value’ Practical Activity: Perform the evaluation of the results obtained from the audit blind sample | Theory- 1 Hrs Practical-4 Hrs Total- 5 Hrs | - Control charts Limits | Lab |

## Module 1021CS-068.	Perform Correctness of Results

Objective: After the completion of this module, the trainee will be able to  perform the correctness of analysis

Duration:  15 Hours			Theory: 3 Hours		Practice: 12 Hours                  Credit Hours:  1.5

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Analyse the Sample | - .Analyse the sample for pH, Electrical conductivity and total Dissolved solids - Analyse the sample for major Anions and Cations. - Calculate the TDS by sum of the concentration of all the constituents in mg/l - Calculate the electrical conductivity as per the standard method. | - Enlist the parameter required to be analyzed - Describe the calculation of Total dissolved solids (TDS) Practical Activity: Analyze the sample for major anions and cations | Theory- 1.5Hrs Practical- 03 Hrs Total- 4.5 Hrs | - Water sample - Autopipett - Standard operation procedure for the analysis of the major anion cations - Instrument instruction manual. | Lab |
| LU2. Perform Anion- cation balance | - Perform the anion-cation balance by the following formula: % difference=100x∑Cations-∑Anions ∑Cations + ∑Anions - The anion-cation sum expressed as milli equivalent per litter must balance - The acceptance criteria is as fallows Anion Sum (meq/L) Acceptable difference 0-3.0 + 0.2meq/L 3.0-10.0 + 0.2% 10.0-800 + 5% | - Describe the basic principle of anion, cation balance - Describe the percent difference calculation formula - Define milli equivalent - Describe the acceptance criteria of anion cation balance Practical Activity: Calculate the % difference of the given data (Anions and Cations) | Theory- 0.5 Hrs Practical-3 Hrs Total- 3.5 Hrs | - Calculator | Lab |
| LU3. Check the measured electrical conductivity (EC) and ions sum | - Measure the electrical conductivity of the sample - .sum the anions and cations - Both the anions and cations sum should be approximately 1/100 of the measured EC value - P4. If either sum does not meet this criteria, re analyse the sample - The acceptance criteria is as fallows 100xanions(or cations) sum, meq/L= (0.9 – 1.1) EC | - Describe the criteria of the anion cation sum and electrical conductivity value Practical Activity: Calculate the measured EC and ion sum | Theory- 0.5 Hrs Practical-3 Hrs Total- 3.5 Hrs | - EC meter | Lab |
| LU4 Check Measured TDS to EC ratio | - Measure the electrical conductivity of the sample - Measure the TDS of the sample - Calculate the ratio of the TDS and EC as Measured TDS/conductivity= 0.55 – 0.7 - The measured TDS to EC ratio should be between 0.55 and 0.7 - If it is outside these limits thee either measured TDS or measure conductivity is suspect; reanalyse the sample | - Describe the parameters required to determine for measuring TDS and EC ratio - Describe the acceptance criteria for TDS/conductivity Practical Activity: Calculate ratio of measured TDS to EC for the given sample | Theory- 0.5 Hrs Practical-3 Hrs Total- 3.5 Hrs | • EC meter |  |

## Module 1021CS-069.	Determine the Method Detection Limits

Objective: After the completion of this module, the trainee will be able to  determine the method detection limits

Duration:  15 Hours			Theory: 3 Hours		Practice: 12 Hours                  Credit Hours:  1.5

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1 Prepare Standard Solutions | . - Prepare standard stock solution in order to calculate MDL prepare solution of concentration approximately four times the estimated limit. - Prepare desired standard concentration solution from the stock solution. | - Define method detection limit - Describe the basic purpose of method detection limit determination Practical Activity: Prepare standard solutions for MDL | Theory- 1Hrs Practical- 04 Hrs Total- 5 Hrs | - Autopipett - Standards/CRM - Instrument instruction manual. - Glassware | Lab |
| LU2. Analyse the Standard Solution and Calculate the Method Detection Limit | - Set the instrument according to equipment instructions - .Analyse the prepared solution 10 times over the course of several days. - Calculate the standard deviation by the usual method. - Calculate the MDL by substituting standard deviation in the following equation: Method Detection Limit (MDL) = t value x Std. deviation Where t = degree of freedom (n-1) | - Describe the required number of observations for the MDL determination - Describe the method detection limit formula Practical Activity: Calculate the method detection limit of given parameter | Theory- 1 Hrs Practical-3 Hrs Total- 4 Hrs | • Autopipett • Standards/CRM • Instrument instruction manual. • Glassware | Lab |
| LU3. Evaluate the MDL results | - If the MDL is much higher than the estimated detection limit, it may indicate problems with the procedure or with the analytical equipment, improperly made reagents and expiration dates for chemicals. - .Re prepare all the standards - Re analyse the standard and calculate the method detection limit. | - Describe the possible reasons if the MDL is higher than the estimated detection limit - Describe and propose the corrective action if MDL is higher than the estimated limit Practical Activity: Calculate the MDL of any parameter | Theory- 1 Hrs Practical-4 Hrs Total- 5 Hrs | - Notepad - Required equipment - Calculator | Lab |

## Module 1021CS-070.	Develop Control Charts

Objective: After the completion of this module, the trainee will be able to  develop the control charts to ensure the accuracy of results

Duration:  15 Hours			Theory: 3 Hours		Practice: 12 Hours                  Credit Hours:  1.5

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Prepare Lab Control Sample for -Control Chart | . - Considering the measurement range of various parameter/methods Prepare lab control samples by following steps - Prepare Stock solution from certified reference material/in-house standards. Desired standard concentration is prepared from the stock solution by using the following formula C1V1 = C2V2 C1 = Initial Concentration V1 = Required Volume C2 = Concentration Required V2 = Final Volume | - Define the basic purpose of control charting - Describe the preparation of lab control sample for developing control chart Practical Activity: - Prepare lab control sample for developing control chart | Theory- 1.5Hrs Practical- 03 Hrs Total- 4.5Hrs | - Autopipett - Standards/CRM - Instrument instruction manual. - Glassware | Lab |
| LU2 Analyse the Lab Control Sample | - Analyse the Lab control sample of each parameter 30 times by different analysts on different days. - .Calculate the Mean as ∑ of results No. of Observations P3: calculate the Standard Deviation (SD) as Xi= No of Observations x¯=Mean Value N= No of readings | - Define standard deviation - Describe the formula of mean and standard deviation Practical Activity: Calculate the standard deviation of given data | Theory- 0.5 Hrs Practical-3 Hrs Total- 3.5Hrs | • Autopipett • Standards/CRM • Instrument instruction manual. • Glassware | Lab |
| LU3. Set the Control Limits | Set the Control Limits as fallowing - Upper control limit UCL = x¯ + 3 s - Upper warning limit UWL = x + 2 s - Lower warning limit LWL = x - 2 s - Lower control limit LCL = x - 3 s | - Enlist three limits of control chart - Describe the formulae for setting control chart Practical Activity: Generate the control chart for given data | Theory- 0.5 Hrs Practical-3 Hrs Total- 3.5Hrs |  | Lab |
| LU4. Interpret the Control chart data | - Control Limits: If one measurement exceeds the CL repeat the analyses immediately. If the repeat measurement is within the CL, continue the analyses. If it exceeds the CL, discontinue the analyses, identify the root cause and correct the problem. - Warning Limits: If two out of three successive points exceeds the WL, analyse another sample. If the next point is within the WL, continue analyses, if the next point exceeds the WL, evaluate potential bias and correct the problem. - Standard Deviation: If four out of five successive points exceeds 1s, or are in decreasing or in increasing order indicate the mean shift. Analyse another lab control sample for confirmation. If the next point is less than 1s, or changes the order, continue analyses, otherwise discontinue analyses and correct the problem. - Trending: If seven successive points are on the same side of the central line, indicate that there has been a shift in the mean; work can continue, however the cause of the shift needs to be identified and the limits re-set if justified | - Describe the various rules for control chart - Describe trend analysis of control chart Practical Activity Perform trend analysis for given control chart | Theory- 0.5 Hrs Practical-3 Hrs Total- 3.5Hrs |  |  |

## Module: 1021 CS-071 Identify Water Resources of Pakistan.

Objective: After the completion of this module, the trainee will be able to develop skill and competence required to identify major water resources of Pakistan

Duration: 40 Hours			Theory:  8 Hours		Practice:  32 Hours                  Credit Hours:  4.0

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Demonstrate water and land potential of Pakistan LU2. Demonstrate irrigation and drainage system LU3. Demonstrate water conservation technologies and practices LU4. Demonstrate rainwater harvesting techniques | Trainee will be able to: - Identify physiographic features - Identify climate and meteorology - Differentiate various kind of geology and soils - Identify land resources including irrigated areas, rain-fed areas, hill torrent areas, desert areas, and wastelands - Determine water resources - Identify Irrigation System of Pakistan - Understand Historic Irrigated Agriculture - Classify Development after Independence - Recognize Indus Basin Irrigation System (IBIS) - Recognize Irrigation Systems Outside Indus Plains including: Rod Kohi/Bandat (Flood Irrigation) - Medium/Small Dams - Sailaba (Falling Flood) Irrigation. - Demonstrate technologies for Irrigated Areas - Demonstrate Watercourse Improvement - Demonstrate Improved Farm Layout - Demonstrate Laser Land Levelling - Demonstrate Improved Irrigation Application Methods - Demonstrate Skimming Wells - Demonstrate Use of Low Quality Groundwater for Agriculture - Demonstrate Irrigation Scheduling - Demonstrate Zero Tillage Technology - Demonstrate Technologies for Barani (rain fed) Areas - Small/Mini Dams - Dug Wells - Demonstrate Pressurized Irrigation System - Demonstrate Technologies for Mountainous Areas - Northern Areas - Demonstrate various types of rainwater harvesting techniques - Prepare design of RWH - Prepare cost estimates for RWH Systems - Implement RWH systems | - Discuss landscape of Pakistan on basis of water resources and topographic features Practical Activity: Demarcate agro-ecological zones of Pakistan on given map Understand irrigation system of Pakistan with various features Practical Activity Draw Indus Basin Irrigation System on a piece of paper and mark various canals and other water bodies Knowledge about various water conservation technologies of various regions Practical Activity Prepare a layout of Drip irrigation system for 1 acre citrus orchard Learn rainwater harvesting system concepts and methods Practical Activity Design a Rainwater harvesting system for a 5 kanals building | Theory- 2 Hrs Practical- 4 Hrs Total- 6Hrs Theory- 2 Hrs Practical- 4 Hrs Total- 6Hrs Theory- 2 Hrs Practical- 4 Hrs Total- 6Hrs Theory- 2 Hrs Practical- 4 Hrs Total- 6Hrs | - Powerpoint presentations - Reports - Books - Powerpoint presentations - Reports Books - Powerpoint presentations - Reports Books - Powerpoint presentations - Reports Books | Class Room Class Room Class Room Class Room |

## Module: Manage meetings with customers

Objective: After the completion of this competency standard, the Trainee will be able to manage a range of meetings including
overseeing the meeting preparation processes, chairing meetings, organizing the minutes and reporting meeting outcomes. It applies to individuals employed in a range of work environments who are required to organize and manage meetings within their workplace, including conducting or managing administrative tasks in providing agendas and meeting material. They may work as senior administrative staff or may be individuals with responsibility for conducting and chairing meetings in the workplace

Duration:  5 Hours			Theory:  5 Hours			Practice:  0 Hours          	        Credit Hours:  0.5

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Plan and Prepare for meetings | Trainee will be able to: - Develop agenda in line with stated meeting purpose - Ensure style and structure of meeting are appropriate to its purpose - Identify meeting participants and notify them in accordance with organizational procedures - Confirm meeting arrangements in accordance with requirements of meeting - Dispatch meeting papers to participants within designated timelines | - Outline responsibilities of the chairperson and explain group dynamics in relation to managing meetings | Theory- 1 Hrs Practical- Hrs Total- 1 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |
| LU2. Conduct meetings | Trainee will be able to: - Chair meetings in accordance with organizational requirements, agreed conventions for type of meeting and legal and ethical requirements - Conduct meetings to ensure they are focused, time efficient and achieve the required outcomes - Ensure meeting facilitation enables participation, discussion, problem solving and resolution of issues - Brief minute-taker on method for recording meeting notes in accordance with organizational requirements and conventions for type of meeting | - Describe options for meetings including face-to-face, tele-conferencing, web-conferencing and using webcams | Theory- 2 Hrs Practical- Hrs Total- 2 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |
| LU3. Follow up meeting’s agenda for implementation | Trainee will be able to: - Check transcribed meeting notes to ensure they reflect a true and accurate record of the meeting and are formatted in accordance with organizational procedures and meeting conventions - Distribute and store minutes and other follow-up documentation within designated timelines, and according to organizational requirements - Report outcomes of meetings as required, within designated timelines | - Devise meeting terminology, structures, arrangements as per organizational procedures | Theory- 2 Hrs Practical- Hrs Total- 2 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |

## Module: Manage workforce planning

Objective: After the completion of this competency standard, the Trainee will be able to manage planning in relation to an
organization’s workforce including researching requirements, developing objectives and strategies, implementing initiatives and monitoring and evaluating trends. It applies to individuals who are human resource managers or staff members with a role in a policy or planning unit that focuses on workforce planning.

Duration:  5 Hours			Theory:  5 Hours			Practice:  0 Hours          	        Credit Hours:  0.5

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Research workforce requirements | Trainee will be able to: - Review current data on staff turnover and demographics - Assess factors that may affect workforce supply - Establish the organization’s requirements for a skilled and diverse workforce | Outline industrial relations relevant to the specific industry | Theory- 1 Hrs Practical- Hrs Total- 1 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |
| LU2. Develop workforce objectives and strategies | Trainee will be able to: - Review organizational strategy and establish aligned objectives for modification or retention of the workforce - Consider strategies to address unacceptable staff turnover, if required - Define objectives to retain required skilled labor - Define objectives for workforce diversity and cross-cultural management - Define strategies to source skilled labor - Communicate objectives and rationale to relevant stakeholders - Obtain agreement and endorsement for objectives and establish targets - Develop contingency plans to cope with extreme situations | Explain current information about external labor supply relevant to the specific industry or skill requirements of the organization | Theory- 1 Hrs Practical- Hrs Total- 1 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |
| LU3. Implement initiatives to support workforce planning objectives | Trainee will be able to: - Implement action to support agreed objectives for recruitment, training, redeployment and redundancy - Develop and implement strategies to assist workforce to deal with organizational change - Develop and implement strategies to assist in meeting the organization’s workforce diversity goals - Implement succession planning system to ensure desirable workers are developed and retained - Implement programs to ensure workplace is an employer of choice | - Strategies to assist workforce to deal with organizational change & work environment | Theory- 1 Hrs Practical- Hrs Total- 1 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |
| LU4. Monitor and evaluate workforce trends | Trainee will be able to: - Review workforce plan against patterns in exiting employee and workforce changes - Monitor labor supply trends for areas of over- or under-supply in the external environment - Monitor effects of labor trends on demand for labor - Survey organizational climate to gauge worker satisfaction - Refine objectives and strategies in response to internal and external changes and make recommendations in response to global trends and incidents - Regularly review government policy on labor demand and supply - Evaluate effectiveness of change processes against agreed objectives | Describe labor force analysis and forecasting techniques | Theory- 2 Hrs Practical- Hrs Total- 2 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |

## Module: Undertake project work

Objective: After the completion of this competency standard, the Trainee will be able to undertake a straightforward project or
a section of a larger project. It covers developing a project plan, administering and monitoring the project, finalizing the project and reviewing the project to identify lessons learned for application to future projects. This unit applies to individuals who play a significant role in ensuring a project meets timelines, quality standards, budgetary limits and other requirements set for the project.

Duration:  5 Hours			Theory:  5 Hours			Practice:  0 Hours          	        Credit Hours:  0.5

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Define project | Trainee will be able to: - Access project scope and other relevant documentation - Define project stakeholders - Seek clarification from delegating authority of issues related to project and project parameters - Identify limits of own responsibility and reporting requirements - Clarify relationship of project to other projects and to the organization’s objectives - Determine and access available resources to undertake project | Give examples of project management tools and how they contribute to a project Explain the organization’s procedures and processes that are relevant to managing a project including: lines of authority and approvals quality assurance human resources budgets and finance recordkeeping reporting | Theory- 1 Hrs Practical- Hrs Total- 1 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |
| LU2. Develop project plan | Trainee will be able to: - Develop project plan in line with the project parameters - Identify and access appropriate project management tools - Formulate risk management plan for project, including Work Health and Safety (WHS) - Develop and approve project budget - Consult team members and take their views into account in planning the project - Finalize project plan and gain necessary approvals to commence project according to documented plan | - Outline types of documents and other sources of information commonly used in defining the parameters of a project | Theory- 1 Hrs Practical- Hrs Total- 1 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |
| LU3. Administer and monitor project | Trainee will be able to: - Take action to ensure project team members are clear about their responsibilities and the project requirements - Provide support for project team members, especially with regard to specific needs, to ensure that the quality of the expected outcomes of the project and document time lines are met - Establish and maintain required recordkeeping systems throughout the project - Implement and monitor plans for managing project finances, resources and quality - Complete and forward project reports as required to stakeholders - Undertake risk management as required to ensure project outcomes are met - Achieve project deliverables | - Outline the organization’s mission, goals, objectives and operations and how the project relates to them | Theory- 1 Hrs Practical- Hrs Total- 1 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |
| LU4. Finalize project | Trainee will be able to: - Complete financial recordkeeping associated with project and check for accuracy - Ensure transition of staff involved in project to new roles or reassignment to previous roles - Complete project documentation and obtain necessary sign-offs for concluding project | - Explain processes for identifying and managing risk in a project | Theory- 2 Hrs Practical- Hrs Total- 2 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |

## Module: Identify and communicate trends in career development

Objective: After the completion of this competency standard, the Trainee will be able to conduct research to identify and communicate career trends. It establishes the need to interact professionally with others in assessing career needs, to effectively assist clients identify competencies

## Module: Apply interpersonal skills

Objective: After the completion of this competency standard, the Trainee will be able to use advanced and specialized communication skills in the client-counselor relationship. This unit applies to individuals whose job role involves working with clients on personal and psychological issues within established policies, procedures and guidelines.

Duration:  5 Hours			Theory:  5 Hours			Practice:  0 Hours          	        Credit Hours:  0.5

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Communicate effectively | Trainee will be able to: - Identify communication barriers and use strategies to overcome these barriers in the client-counselor relationship - Facilitate the client-counselor relationship through selection and use of micro skills - Integrate the principles of effective communication into work practices - Observe and respond to non-verbal communication cues - Consider and respond to the impacts of different communication techniques on the client-counselor relationship in the context of individual clients - Integrate case note taking with minimum distraction | Legal and ethical considerations for communication in counseling practice, and how these are applied in individual practice: codes of conduct/practice discrimination duty of care human rights practitioner/client boundaries privacy, confidentiality and disclosure rights and responsibilities of workers, employers and clients work role boundaries responsibilities and limitations of the counselor role work health and safety Principles of person-centered practice Communication techniques and micro-skills including: attending behaviors active listening, reflection of content feeling, summarizing questioning skills open, closed, simple and compound questions client observation skills noting and reflecting skills providing client feedback Specialized counseling communication techniques, and how they are used, including: challenging reframing focusing K8: components of the communication process including: encoder decoder Obstacles to the counseling process Self-evaluation practices, including: how to recognize own biases Impact of own values on the counseling relationship. | Theory- 1 Hrs Practical- Hrs Total- 1 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |
| LU2. Use specialized counseling interviewing skills | Trainee will be able to: - Select and use communication skills according to the sequence of a counseling interview - Identify points at which specialized counseling interviewing skills are appropriate for inclusion - Use specialized counseling communication techniques based on their impacts and potential to enhance client development and growth - Identify and respond appropriately to strong client emotional reactions | - Key objectives of counseling interview - Stages of a counseling interview - Potential impacts of using different communication skills and techniques in counseling contexts | Theory- 2 Hrs Practical- Hrs Total- 2 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |
| LU3. Evaluate own communication | Trainee will be able to: - Reflect on and evaluate own communication with clients - Recognize the effect of own values and beliefs on communication with clients - Identify and respond to the need for development of own skills and knowledge | Impacts of trauma and stress on the communication process, including on: concentration and attention memory use of verbal and written language use of body language challenging within the counseling session Observational techniques including: facial expressions non-verbal behavior posture silence Ways in which different people absorb information, including: visual auditory kinesthetic K10: communication barriers and resolution strategies, including: environmental physical individual perceptions cultural issues language age issues disability K9: primary factors that impact on the communication process including: context participants rules messages noise feedback | Theory- 2 Hrs Practical- Hrs Total- 2 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |

Module: Work safely in an office environment

Objective: After the completion of this competency standard, the Trainee will be able to participate in workplace occupational health and safety (OHS) processes to protect workers own health and safety, and that of others.

Duration:  5 Hours			Theory:  5 Hours			Practice:  0 Hours          	        Credit Hours:  0.5

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Work safely | Trainee will be able to: - Follow established safety procedures when conducting work - Carry out pre-start systems and equipment checks in accordance with workplace procedures | - Explain responsibilities of employers and employees under relevant health and safety regulation | Theory- 1 Hrs Practical- Hrs Total- 1 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |
| LU2. Implement workplace safety requirement | Trainee will be able to: - Identify designated persons for reporting queries and concerns about safety in the workplace - Identify existing and potential hazards in the workplace, report them to designated persons and record them in accordance with workplace procedures - Identify and implement workplace procedures and work instructions for controlling risks - Report emergency incidents and injuries to designated persons | - Report of potential hazards at the workplace to designated persons and record at workplace - Workplace procedures and work instructions for controlling risks | Theory- 1 Hrs Practical- Hrs Total- 1 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector - Safety signs - Safety procedures - Variety of forms, formats used within the organization | Class Room |
| LU3. Participate in OHS consultative processes | Trainee will be able to: - Contribute to workplace meetings, inspections or other consultative activities - Raise OHS issues with designated persons in accordance with organizational procedures - Take actions to eliminate workplace hazards or to reduce risks | - Describe emergency procedures including procedures for fires, accidents and evacuation | Theory- 1 Hrs Practical- Hrs Total- 1 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |
| LU4. Follow safety procedures | Trainee will be able to: - Identify and report emergency incidents - Follow organizational procedures for responding to emergency incidents | - Outline commonly used hazard signs and safety symbols | Theory- 2 Hrs Practical- Hrs Total- 2 Hrs | - Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector - Safety signs - Safety procedures - Variety of forms, formats used within the organization | Class Room |

## Module: Maintain professionalism in the workplace

Objective: After the completion of this competency standard, the Trainee will be able to maintain a professional image in the
workplace, including behaving ethically, demonstrating motivation, respecting timeframes and maintaining personal appearance.

Duration:  5 Hours			Theory:  5 Hours			Practice:  0 Hours          	        Credit Hours:  0.5

| Learning Unit | Learning Outcomes | Learning Elements | Duration | Materials Required | Learning Place |
| --- | --- | --- | --- | --- | --- |
| LU1. Respect work timeframes | Trainee will be able to: - Demonstrate punctuality in meeting, set working hours and times. - Utilize working hours only for working and follow company regulations. - Complete work tasks within deadlines according to order of priority - Supervisors are informed of any potential delays in work times or projects | - Explain application of good manners and right conduct - Outline the company code of conduct/values | Theory- 1 Hrs Practical- Hrs Total- 1 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |
| LU2. Maintain personal appearance and hygiene | Trainee will be able to: - Clean hair, body and nails regularly. - Wear suitable cloths for the workplace, and respect local and cultural contexts - Meet specific company dress code requirements | - Describe common products used for oral and personal hygiene | Theory- 1 Hrs Practical- Hrs Total- 1 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |
| LU3. Maintain adequate distance with colleagues and clients | Trainee will be able to: - Respect personal space of colleagues and clients with reference to local customs and cultural contexts. - Keep sufficient distance from others - Avoid cross transmission of infections (especially through respiration) | - State workplace hygiene standards - Explain basic practices for oral and personal hygiene | Theory- 1 Hrs Practical- Hrs Total- 1 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Mulmedia projector | Class Room |
| LU4. Work in an ethical manner | Trainee will be able to: - Follow company values/ethics codes of ethics and/or conduct, policies and guidelines. - Use company resources in accordance with company ethical standards. - Conduct personal behavior and relationships in accord with ethical standards and company policies. - Undertake work practices in compliance with company ethical standards, organizational policy and guidelines. - Instruct co-workers on ethical, lawful and reasonable directives. - Share company values /practices with co-workers using appropriate behavior and language. - Report work incidents /situations and/or resolved in accordance with company protocol /guidelines. | - Outline the Company regulations, performance and ethical standards - Explain work responsibilities/job functions - Describe communication skills | Theory- 2 Hrs Practical- Hrs Total- 2 Hrs | Internet MS Office Smartphone Laptop / Computer Smartphone Printer White board Projector screen Multimedia projector | Class Room |

# List of Tool & Equipment:

As mentioned in each module.